## Sheet1
2010-02-17 |  | NML Injector Laser Lab Room Requirements
FINALIZED
| HVAC | Room static heat | 13KW nom, + 3 people
|  | Temperature | 70oF, dT<1oF (<+/- 0.5oF)
|  | Humidity | 40 +5/- 10%
|  | HEPA filter | Class 10,000 Clean Room:positive pressure; laminar down flow
| Utilities | Electrical Power | power in room is atleast 13KW nom (calc), 33 KW peak (measured)
480VAC, 90A, 75KVA transformer to feed 208Y panel
480VAC, 30A ckt for electric duct heater
42 circuit 208Y VAC; 220 amps panel
Outlets distributed at tables racks and in 36" high wire chase in room perimeter
|  | Water (LCW) | 70oF; 12GPM; 1.0 MΩ, Neslab chiller to exchange with high-bay air or 95LCW
X |  | Dry Air | at least 3 CFM, 80-120psi, dry and oil free, could be dry N2.
X |  | Dry N2 boil-off | 1/4" line of 20-120 psi, for cleaning optics
|  | Network | at least twelve 1Gbit RJ45 spiggots for desktops and instruments
Wi-Fi for laptops and instruments
X |  | Telecom | 1 digital telecom line for Laser Lab phone number. Also need x5699 and x4996.
X | Room | Light Tight | Light should not be allowed to enter or exit the room.
|  | Doors and windows | Door needs to be double 36" metal doors that swing out of room.
X |  |  | No windows. Light tight seals.
X |  | Floor | Epoxy coating on concrete floor, gray in color
6" raised floor (10-12 HVAC return panels).
X |  | Room | water leaks should not be allowed into room.
|  | Walls | 2"x6" architectural studs
X |  |  | walls insulated for sound
X |  |  | 5/8" fire rated drywall on both sides
X |  |  | interior painted flat black to reduce reflections.
|  | Ceiling | Drop ceiling with kitchen grade panels, flat black in color
flush floresent lighting pannels, three switch legs
X | ES&H | Interlocks | Room interlocked as an "Enclosure" with Crash buttons and Fire key
X |  |  | Store-Room lock on door with core controlled by AD/ES&H.
X |  |  | Shutter permit; all laser shutters monitored and fail closed.
X |  |  | Power Supply Permit; Crash Buttons, Fire Key, room interlock, shutter failure mode
X |  |  | Transport Permit; cave secure, vacuum good, alignment HeNe bypass mode
X |  | Fire Protection | Smoke detectors with external relay curcuit for Safety system
X |  |  | Halon fire extinguisher (if available), CO2 extinguisher
X |  |  | VESDA system would be useful if money allows
Sprinkler system should be Halon or dry. Water leaks not tolerated

## Sheet2
The room should be located and oriented such that the RF is located as close as possible to the last UV mirror in laser room. The length of UV transport line should not be longer than 20 m.
2.Laser room should not be far from the control room, as various signal cables have to be laid between two rooms.
Vibration of the floor should be sub-wavelength (1053 nm). Preferably the integrated vibration < 20 nm.
The temperature fluctuation should be within 1°F (.<± 0.5 °F). It is preferred to keep positive pressure in laser room so that the air flow will be stable and uniform.
Humidity should be 45% ± 5%.
UV transport line has to be vacuumed and the boxes at the turning corners are needed to hold the mirrors and lenses.
Cleanliness in laser room should be class 10,000 or better.
Cooling water 1MOhm, 70 °F,12 gallon/min. With the upgrade of diode pump amplifiers this requirement could be eased a lot.
Power requirement (208 VAC, 200 A).
It will be nice if we can have some platforms and cable trays over the optical tables hanging from ceilings so that we can put staffs on them rather than on the towers being built around the tables.
Need walls and ceiling smooth, painted, and clean.
The floor should be epoxy coated for clean room.
Miscellaneous (e.g. fire alarm and interlocked smoke detector, CO2 fire extinguisher).