## Slide 1: TE1AES004 JLAB T-map test
G. Ciovati, M. Morrone, G. Wu

## Slide 2
Cavity sealed on 6/7,
T-map installed on 6/8,
RF tested on 6/9,
Disassembled on 6/10,
Shipped back on 6/11.
Cavity see no chemical etching since 2008,
Only see HPR and molding.
Cavity last tested at A0 (40 MV/m)
No active pumping since A0 RF test.
No x-ray was detected at JLAB (38 MV/m)
Three hot spots and one quench site located
Geometric defect showed no heating at all
Previous quench site (A0) was a hot spot
Quench had pre-heating

## Slide 3
Defect
Previous quench site

## Slide 4
Defect
Previous quench site

## Slide 5
Right before MP
Defect
Previous quench site

## Slide 6
Right after MP
Defect
Previous quench site

## Slide 7
Defect
Previous quench site

## Slide 8
31.63 MV/m
Defect
Previous quench site

## Slide 9
34.86 MV/m
Defect
Previous quench site

## Slide 10
36.83 MV/m
Defect
Previous quench site

## Slide 11
37.56 MV/m
Defect
Previous quench site

## Slide 12
37.72 MV/m
Defect
Previous quench site

## Slide 13
37.85 MV/m
Defect
Previous quench site

## Slide 14
36.75 MV/m
Defect
Previous quench site

## Slide 15
36.75 MV/m
Defect
Previous quench site

## Slide 16
Right after quench
Defect
Previous quench site
Quench at new location

## Slide 17
Right after quench

## Slide 18
No x-ray was detected at JLAB (38 MV/m)
Three hot spots and one quench site located
Geometric defect showed no heating at all
Previous quench site (A0) was a hot spot
Quench had pre-heating
Summary
Why is the quench site changed test after test?
- Cut out may tell us why.