## Slide 1: Status of 1.3-GHz Cavity Vacuum Furnace for IB4
Last week’s testing
  4 hours pumping - 600°C
  Another 8-10 hours pumping - 1000°C
  Baked at 1100°C for 2 hours to clean chamber
  Cooled but continued pumping to reach ultimate vacuum pressure over weekend

## Slide 2: Vacuum Pressure Measurements
Chamber temperature (°C) | Cold Cathode reading (torr) | RGA reading (torr) | Required by Spec (torr)
70 | -- | 1 x 10-7 | 1 x 10-8
600 | 9.9 x 10-8 | 9 x 10-7 | 1 x 10-7
1000 | 1.8 x 10-7 | 2 x 10-6 | 1 x 10-6

## Slide 3: Current Schedule
Week of July 26: Acceptance testing at vendor site
Week of Aug 2: Preparation for shipping
Week of Aug 9: Shipping and delivery to FNAL
Aug 16 – Sept 20:
  Installation
  Final acceptance testing
  Operator training @ FNAL
  ORC Review