## Slide 1: The Sample Test of the FNAL/IB4 Vacuum Furnace for Hydrogen Degassing
G. Wu

## Slide 2: Niobium RRR and impurities

## Slide 3: RRR Niobium Samples
Niobium samples are cut from cavity production sheet
  Low impurities of H.O.N.C.
  Annealed with low crystal defects
Niobium sample preparation
  Measure RRR
  BCP etching or ultrasonic clean
  Bake out in furnace
  Measure RRR

## Slide 4: Furnace vacuum quality indicators
O, N, C, H are the main residuals in a furnace; Niobium is a perfect getter, O is the prime source among them that can change the niobium RRR.
Chart: A typical RGA data
Furnace vacuum quality is essential. Indicators are total pressure, oxygen pressure and hydrogen pumping capacity

## Slide 5: RGA reading during baking on 10/19/2010
Oxygen partial pressure was inaccurate. Vacuum lost when samples were still hot (400-700C).
Air supply stopped

## Slide 6: RGA reading during baking on 10/26/2010
Oxygen partial pressure was nicely low, averaged in 1x10-9 Torr

## Slide 7: Three tests of RRR niobium samples
RRR improved in third test indicating furnace is compatible to niobium
With Stainless Steel Tube, no RGA logging
With Stainless Steel Tube, see RGA logging
No Stainless Steel Tube, see RGA logging

## Slide 8: TE1ACC004/TE1CAT002 baking

## Slide 9: RGA reading during baking (600 C) at ANL/APS
Compared to ANL/APS furnace

## Slide 10: RGA reading during baking (600C) at JLAB
Compared to JLAB furnace

## Slide 11: RGA reading during baking (800C) at JLAB
Compared to JLAB furnace
Temperature
H
O
9-cell PKU-2

## Slide 12: Partial pressure comparison
Furnaces | H pressure [torr] | O pressure [torr] | Note
JLAB (800C) | 0.4-1.2 x10-5 | <6 x10-9 | No Q-disease
ANL/APS (600C) | 0.4 x10-5 | <0.5 x10-9 | No Q-disease
“Little TIM” (800C) | 0.6 x10-5 | <0.7 x10-9 | ?
H O partial pressure at the end of baking cycle

## Slide 13: Conclusions
The FNAL/IB4 furnace is acceptable for cavity baking
The hydrogen degassing for cavity is effective, to be confirmed by Q-disease measurement of 1-cell cavities.