## Slide 1: Electrical Properties of MPPC/SiPM/GMAPD’s
Adam Para
June 24, 2011

## Slide 2: Motivation
MPPC/SiPM/GMAPD are avalanche diodes biased in the reverse direction and operated in a Geiger mode (so we think)
Their behavior depends on the electrical properties of a cell: resistance and capacitance, hence it is important to measure these properties for different detectors we are trying to understand
Detectors:
  1 mm x 1 mm, 20m pixel 025U, 50m pixel 050U, 100m pixel (100U)
  3 mm x 3 mm, 50m pixel, 050PX

## Slide 3: Ideal Diode (Shockley)
Current in the forward direction, where
thermal voltage
saturation current

## Slide 4: I-V Curve for a Real Detector
We have a quenching resistor in series with the diode

## Slide 5: Hamamatsu 025U
1 mm x 1 mm, 1600 pixels

## Slide 6
the same data in a linear and logarithmic scale
I-V characteristics measured at temperatures from -60C to +50C
Shockley equation provides a very good description of the measured I-V curves over 5 orders of magnitude
slopes of the lines on a linear plot => total resistance
Assuming the parasitic resistance is small, the quenching resistance Rq=NpixelxRdev
slopes of the lines on the log plots => thermal voltage
vertical offset on the logarithmic plot => saturation current
all parameters change with temperature

## Slide 7: I-V Diode Only

## Slide 8: Quenching Resistance
Quenching resistor gets smaller as the temperature rises (polysilicon!)
Non-linear dependence of the quenching resistance on temperature

## Slide 9: Thermal Voltage
fitted values of the thermal voltage change linearly with temperature, as expected
quenlity factor n very close to 1 (1.04)

## Slide 10: Saturation Current
Saturation current varies by 10 orders of magnitude.
Presumably related to dark pulses rates

## Slide 11: Hamamatsu 050U
1 mm x 1 mm, 400 pixels

## Slide 14: Hamamatsu 100U
1 mm x 1 mm, 100 pixels

## Slide 17: Hamamatsu 050PX
3 mm x3 mm, 3600 pixels

## Slide 20: Comparison of detectors

## Slide 21: Quenching Resistance
50U and 100U have very similar quenching resistor
25U has about twice as big
50PX has the resistor 50% higher then 50U

## Slide 22: Thermal Voltage
25U, 50U and 100U have ‘the same diode’
50PX diode seems to be somewhat different

## Slide 23: Saturation Current
25U, 50U and 100U have ‘the same diode’
50PX diode seems to be somewhat different