## Slide 1: Design of the Timepix3
X. Llopart
7th July 2011
LCTPC collaboration meeting

## Slide 2: Outline
Introduction to Timepix3
Timepix3 Analog Front-End
Timepix3 readout architecture
Conclusions

## Slide 3: 2 Collaborations → 21 collaborating institutes
Medipix2
INFN Cagliari
CEA-LIST Saclay
CERN Geneva
University of Erlangen
University of Freiburg
ESRF Grenoble
University of Glasgow
University of Houston
IFAE Barcelona
Mid Sweden University
MRC-LMB Cambridge
INFN Napoli
NIKHEF Amsterdam
INFN Pisa
FZU CAS Prague
IEAP CTU Prague
SSL Berkeley
Medipix3
ALMOF Amsterdam
University of Bogota
University of Canterbury NZ
CEA-LIST Saclay
CERN Geneva
DESY Hamburg
Diamond Light Source
University of Erlangen
ESRF Grenoble
University of Freiburg
University of Glasgow
ITER
University of Karlsruhe
Leiden University
Mid Sweden University
NIKHEF Amsterdam
IEAP CTU Prague
SSL Berkeley
VTT Microsystems
http://www.cern.ch/medipix

## Slide 4: The Medipix Chips
A philosophy of functionality built into the pixel matrix allows complex behavior with a minimal inactive region
55um square pixel matrix
256 by 256
Configurable ‘shutter’ allows many different applications
Silicon, 3D, CdTe, GaAs, Amorphous Silicon, Gas Amplification, Microchannel Plates etc…

## Slide 5: Development Path
Medipix1
1um SCAMOS 64 by 64 pixels
Photon Counting Demonstrator (1997)
250nm IBM CMOS, 256 by 256 55um pixels
Full photon counting (2002)
130nm IBM CMOS, 256 by 256 55um pixels
Photon Counting, Spectroscopic, Charge Summing, Continuous Readout (2009)
Fast front end, Simultaneous ToT and ToA (2012)
<20 Mcounts/cm2/s
Analogue (ToT) and Time Stamping (ToA) (2006)
Medipix2
Timepix
Medipix3
Timepix3
VELOpix
CLICpix
65nm ? , ~20um pixels
Future Hybrid Pixel Time tagging layer for the LCD project (20??)
Future LHCb readout – Data driven 40MHz ToT 12Gb/s per chip (2013)
~500 Mcounts/cm2/s
Slide from R.Plackett (Diamond LS)

## Slide 6: Timepix3 Scope
Several groups in the Medipix3 collaboration have shown interested in a new version of the Timepix (2006) → Timepix3
Large range of applications (HEP and non-HEP):
  X-ray radiography, X-ray polarimetry, low energy electron microscopy
  Radiation and beam monitors, dosimetry
  3D gas detectors, neutrons, fission products
  Gas detector, Compton camera, gamma polarization camera, fast neutron camera, ion/MIP telescope, nuclear fission, astrophysics
  Imaging in neutron activation analysis, gamma polarization imaging based on Compton effect
  Neutrino physics
  Main Linear Collider application: pixelized TPC readout
Reuse many building blocks from Medipix3 chip (2009)
Timepix3 is an approved project by the Medipix3 collaboration with an assigned budget (2-engineering runs)
Design groups: NIKHEF, BONN, CERN

## Slide 7: Required readout chips for the HEP applications
Upgrade of the LHCb VELO detector: VELOPIX
  Minimum pixel size
  Extremely high illumination:
    up to 290 Mhits/cm2/s
  Continuous data readout
  Sparse data (zero suppression)
  Charge measurement
  Micro-pattern gas avalanche detectors: GOSSIP, TPC
  Minimum pixel size
  High time resolution: ~ 1ns
  Wide dynamic range : ~ 100μs
  Charge measurement
  High sensitivity : threshold ≈ 350e-
Slide from V.Gromov (Nikhef)

## Slide 8: Timepix3 Main Requirements
Matrix layout: 256x256 pixels (Pixel size 55x55 µm)
Time stamp and TOT recorded simultaneously
  10 bit Energy Measurement (TOT)
    Standard Resolution 25ns (@40MHz)
    Energy Dynamic range up to 25.6 µs (@40MHz)
  14 bits Slow time-stamp
    Resolution 25ns (@40MHz)
    Dynamic range 409.6 µs (14 bit)
  4 bits Fast time-stamp
    resolution ~1.6ns (using on-pixel oscillator running at 640MHz)
    Dynamic range 25ns
Sparse Readout
Technology choice: IBM 130nm DM4-1 (as in Medipix3)

## Slide 9: The Timepix3 Chip
This chip will be
220 µm
110 µm
Readout Chip | TIMEPIX3 (beginning of 2012)
Pixel size | 55 x 55 µm2
Pixel arrangement | 256 x 256 (2x4 superpixels)
Sparse readout | YES
Operating Modes | 1: TOT & TOA 2: TOA 3: Event Counting + TOT
Minimum threshold | > 500 e- (1.8 keV)
TOA resolution | > 1.5 ns
Peaking time | < 25 ns
TOT resolution | < 5% channel to channel spread
Technology | IBM 130nm DM 3-2-3 or DM 4-1-3
Power consumption ON | <700 mW (~20 μW/pixel) @1.2 V
Power consumption OFF | <10 mW (Power Pulsing)
Target floorplan | 3 sides buttable and minimum periphery
TSVs possibility | YES. Multi-dicing scheme as Medipix3
Count Rate | 20x 106 particles/cm2/s
Analog
Digital
Analog

## Slide 10: Timepix3: top level block diagram
128 double-columns @ 256● 55μm = 1.4cm
double column data bus (40MHz)
DAC
Ring oscillator (640MHz)
Time Stamp 14b
ToT Counter 10b
Fast Counter 4b
DAC
MUX
Readout
Control (TOKEN based)
Analog Front-end
Digital Region
8-Pixel Region
Double column bus
ToT Counter 10b
Fast Counter 4b
End of Column Logic
FiFO
Readout control
(Token based)
periphery bus
8 x LVDS
Serializer / FIFO
Tx
Data output block
periphery data bus 40MHz
Data_out Data_in
Slow control
RX
Tx
PLL
DAC
EFUSE
chip ID
Bandgap
Clock 40 MHz
RX
Reset_GLB
RX
Bias gen.
MUX
MUX
FIFO
Clock
Vcntr
(oscillator)
Time Stamp 14b
Analog regions
(30%)
Digital region
(70%)
Pixel: 55μm x 55μm
Super Pixel: 8pixels @ 4 x 2
64 Super Pixels in a DC
128 Double Columns (DC)
Super Pixel FIFO : 2-events
DC bus: 40MHz
End-of-DC FIFO: 4-events
Periphery bus: 40MHz @ 44b
Output Serializer / FIFO
Slide from V.Gromov (Nikhef)

## Slide 11: Timepix3 Analog front end

## Slide 12: Timepix3 pixel FE
Timepix3 and VeloPix chips can use the same front-end design
Main requirements:
  Bipolar input with DC leakage current compensation
  Small and low power → 55 µm x [10…20] µm and < 10 µW
  TimeWalk < 25ns → Fast peaking time in Preamp and Fast discriminator (power penalty)
  Full chip minimum detectable charge as low as possible (< 500e-) → 4 bits threshold tuning and low Preamp Noise (ENC)
  Precise TOT measurement → Minimize Preamp Noise (ENC)
The full front-end schematic is designed (CERN and NIKHEF)
Simulation based results are shown in following slides

## Slide 13: FE of Timepix3
The design of Timepix3 FE is based in the Krummenacher FE architecture
10 000e-
5 000e-
1 000e-
-1 000e-
-5 000e-
-10 000e-
holes (h+)
electrons (e-)

## Slide 14: Tunable return to zero (Vout vs Ikrum)
Qin=2ke-
Threshold =1Ke-

## Slide 15: Time-over-threshold response
Preamp Output
Qin=2Ke- to 30Ke- @Ikrum=10nA
Discriminator Output
Qin=2Ke- to 30Ke- @Ikrum=10nA
Threshold =1Ke-

## Slide 16: TOT linearity and ΔTOT/TOT
Threshold=1ke-
Channel ΔTOT/TOT ~10% @3Ke-
Channel ΔTOT/TOT ~1.9% @10Ke-
Channel ΔTOT/TOT ~100% @1.35Ke-

## Slide 17: Channel Noise and Minimum detectable Charge
Target = <500e-
Minimum detectable charge
Pixel noise (ENC) vs Cd[fF]

## Slide 18: TimeWalk Simulations
Threshold=1ke-

## Slide 19: Simulations Summary
 | Simulation Results
Pixel size | 55 µm x 55 µm
Analog pixel area | 55 µm x [10…20] µm
Pixel matrix | 256 x 256
Input charge | Bipolar (h+ and e-)
Leakage current compensation | YES
Detector capacitance (Planar detector) | < 50 fF
Peaking time | <25ns if Cd<70fF
Preamp output linear dynamic range | < 15 Ke-
Timewalk | <25ns [@ Qin>Threshold]
Return to zero (Tunable) | YES (ΔIkrum)
TOT linearity and range | Monotonic up to ~150 Ke- @Ikrum 10nA
Return to zero full chip spread (TOT spread) | < 5%
ENC (σENC) | 0.57*Cd[fF]+61 [e-] → ~75 e- @ Cd=25fF
# Thresholds | 1
Discriminator response time | <5ns [if Qin>Threshold + 5 Ke-]
MC simulated Threshold spread | ~200 e-
Threshold spread after tuning (4 bits) | ~20 e-
Full chip minimum detectable charge | 6*sqrt(ENC2+Threshold_mismatch2) → ~475 e- @ Cd=25fF
Pixel analog power consumption | 8.4 µW static (Preamp 3.6 µW and 4.8 µW Discriminator) 2.4 µW dynamic (only @ HIT)

## Slide 20: Timepix3 readout architecture

## Slide 21: Super Pixel and Pixel
Clock-signal has to be distributed into all modules in
the super pixel.
Pixels are token arbitrated, token released when a hit
is present.
1-bit serial interface between pixels and the readout
logic.
Super pixel sends data to EoC by the column bus.
One ToA time stamp counter per super pixel (or global
time stamp bus).
Slide from T.Poikela (CERN)

## Slide 22: Timing Diagram of Digital Pixel
Slide from T.Poikela (CERN)

## Slide 23: Super Pixel-to-EoC data transport
Bus protocol does not allow any data overflows in the periphery. If EoC is full, no data is transmitted.
Protocol requires a counter in the super pixel to count the number of words per bus transaction (max 3-bit counter)
Global signals in Double Column: Data bus 5b/9b (not fixed yet)
        EoC Full 1b
        DataPresent 1b
        RefClk 1b
        Reset 1b
        Shutter 1b
        Time stamp 14b (or implemented with local counter)
        Total 10b – 28b
Slide from T.Poikela (CERN)

## Slide 24: Timepix3: modes of operation
 | TOA & TOT | Only TOA | Event Count & Integral TOT
Data format | Fast Time (640MHz @ 4b) & Slow Time Stamp (40MHz @ 14b) & ToT (40MHz @10b) & Pixel coordinate (16b) | Fast Time (640MHz @ 4b) & Slow Time Stamp (40MHz @ 14b) & Pixel coordinate (16b) | Event Count(10b) & Integral ToT ( 14b) & Pixel coordinate (16b)
Double hit resolution (per pixel) | ToT + 700ns | > 450ns | -
Readout | continuous sparse data readout with zero-suppression |  | non-continuous sparse data readout with zero-suppression
 |  |  | Max counting rate < 100kHz/pixel
 | Max counting rate < 1kHz/pixel |  | 
Full chip readout time | - |  | 1.6msec
Slide from V.Gromov (Nikhef)

## Slide 25: Timepix3 periphery architecture
Two readout modes:
  Encoded Clock – Data output (at 640 MHz)
  Separate Clock and Data outputs (at 320MHz, 160MHz …. )
Selectable number of data links (1 to 8)
Slow control
Phase
Locked
Loop
SyncCLK
(40MHz)
Mode of operation
definition
Configuration data (threshold DAC’s, masks etc)
Periphery Data Bus: 44bit @ 40MHz
8b/10b
Encoder
FIFO
Serializer
44bit
8bit
10bit
Data
1bit
Round-robin Distributor
DataOut [0:7]
DataIn
ReadoutCLK
MUX
640MHz
320MHz
160MHz
SyncCLK
(40MHz)
:10
½ Clock delay
Clock
1bit
Data_serial_ link
Clock_serial_link
Control
Logic
Shutter
Reset
Busy
Link is almost full
Ext. Readout clock
Slide from V.Gromov (Nikhef)

## Slide 26: Conclusions

## Slide 27: Conclusions
The Timepix3 will be the next chip in the Medipix family
  TOT and TOA simultaneously
  Sparse and data driven readout
  Arrival time resolution ~1.6ns
  Hit rate up to 20x 106 particles/cm2/s
Timepix3 design schedule is suffering from lack of manpower due to other commitments (Medipix3 and FEI4)
Submission of the Timepix3 chip is expected by beginning of 2012
27
220 µm
110 µm

## Slide 28
Spare Slides

## Slide 29: A high density and low power Standard Cell Library

## Slide 30: An high density low power 130 nm digital library
The design of the Timepix3 goes in parallel with the Medipix3.2 design
The Medipix3.2 redesign requires significant changes in the digital part of the pixel (designed as a full custom layout)
Can we use CERN default standard cell 130nm library at the pixel level?
  Too big: @CERN the default library has large cells since is targeted for ~800MHz !!
  Too much leakage power: @CERN the default library is not low-power
A custom made high density library (SC_130nm_XL) design has been done with the initial idea of using it in Medipix3.2 pixel but with a huge potential impact in later developments (Timepix3, Velopix…)
Main actions:
  Reduce cells height → We don’t need big buffers (speed)
  Keep all transistor small or minimum size (W/L=0.28/0.12)
  ADD NV and PV layers → Low power transistors (no area penalty)

## Slide 31: Medipix3 pixel counter (24 bits) + control logic
Synthesized using the CERN IBM 130nm Standard Cell Library (CMOS8RF)
Area is too big (52 x 33.6) → ~60% pixel area
High static (leakage) power consumption:
55 μm
VDD | Temp | Leakage in cell | leakage In Chip
1.4 V | 125 C | 22.5 μW | 1.5 W !!!
1.5 V | 25 C | 223 nW | 14.6 mW
1.6 V | -55 C | 470 pW | 30 μW

## Slide 32: SC_130nm_XL Library
Row Height is fixed to 2.4 µm
Well Tap library
Maximum frequency <~350 MHz
Low power transistors
~45 cells are available in the library (growing fast…)
  No SEU registers (so far)
Encounter Library Characterizer (ELC) has been used to fully characterize the library in different corners
  Full Synopsis library
  Verilog library
  LEF files
6 µm
4.8 µm
Minimum DFF available in default CERN Standard Cell 130nm library
5.6 µm
2.4 µm
Minimum DFF In the custom made High density 130nm library
x2 smaller !!!
1.8 µm
3.8 µm
Minimum DFF available in a commercial Standard Cell 65nm library
x4 smaller !!!

## Slide 33: Medipix3 pixel counter (24 bits) + control logic
Synthesized using the SC_130nm_XL library
Area (52 x 16.8) → ~29% pixel area
  Low cell leakage power
55 μm
VDD | Temp | Leakage in cell | leakage In Chip
1.5 V | 25 C | 2.5 nW | 163 μW
1.2 V | 25 C | 1.15 nW | 75.3 μW

## Slide 34: Comparison between old and new counter
Medipix3.0 (2009)
Medipix3.2 (2011) using SC_130nm_XL

## Slide 35: Medipix3.2 counter using SC_130nm_XL library
RX & PC
+ M1
+ M2
110 µm
19.2 µm