## Slide 1: Cavity BPM noise floor
N. Blaskovic

## Slide 2: Outline
N. Blaskovic
Calibrate I’ using Epics positions
  ADC counts / um
Find noise jitter in:
  ADC counts
  um
Generic waveform of I, Q or I’

## Slide 3: I’ calibration using Epics
Single sampling point
Single BPM
Each circle is a trigger
Each colour is an AQD0FF mover setting
N. Blaskovic

## Slide 4
N. Blaskovic
gradient
-38.64 ADC counts/um
gradient
40.87 ADC counts/um

## Slide 5: Noise floor
Looking at
  arbitrary sampling point before bunch arrival
  arbitrary file setting and BPM
Noise: 11-16 ADC counts
Using ~40 ADC counts/um, noise: ~0.3 um
This sets the BPM resolution lower limit (Colin)
N. Blaskovic

## Slide 6: Spares

## Slide 7
N. Blaskovic
I’ calibration depends on the sample number used for calibration…
A calibration close to the peak value was chosen to provide the lowest limit on
the noise estimate in um