## Slide 1: Results of a beam test of a Large Micromegas TPC prototype with 6 modules
LCWS12 Arlington
October 22-26, 2012
D. Attié, P. Colas, M. Dixit, P. Hayman, T. Maerschalk, A. Robichaud, J. Timmermans, W. Wang

## Slide 2: All the R&D is gathered in LCTPC
38(*) member + 7 observer institutes from 12 countries
(*) 25 signed the MOA
www.lctpc.org

## Slide 3: The EUDET test setup at DESY
The EUDET (FP6) setup at DESY is operational since 2008
Upgraded within AIDA (FP7): autonomous magnet with 2 cryo-coolers since July
SiPM trigger
Field cage

## Slide 4: Charge spreading by resistive foil
Resistive coating on top of an insulator: Continuous RC network which spreads the charge from s(avalanche)~15µ to mm: matching pad width improves position sensitivity
M. Dixit, A. Rankin, NIM A 566 (2006) 28
Z=20cm, 200 ns shaping
PAD RESPONSE: Relative fraction of ‘charge’ seen by the pad, vs x(pad)-x(track)
x(pad) – x(track) (mm)

## Slide 5: Integrated electronics
New detector : new routing to adapt to new connectors, lower anode resistivity (3 MW/sq), new res. foil grounding on the edge of the PCB.
New 300 points flat connectors (zero extraction force)
New front end: keep naked AFTER chips and remove double diodes (count on resistive foil to protect against sparks)
New Front End Mezzanine (FEMI)
New back-end for up to 12 modules
New DAQ, 7-module and more compact format
New trigger discriminator and logic (FPGA).
D. Calvet et al., IEEE Real Time 2012

## Slide 6
The multi-module configuration allows new studies :
Performance in cracks
Module misalignment
distortions

## Slide 7
Smooth data taking in July, with over 1 000 000 evts. Air cooling and temperature control. However 6 modules only, 2 of them prototype or pre-serie, imperfect contacts

## Slide 8
Thanks to Bo Li, Keisuke Fujii, Gilles de Lentdecker, M. Killenberg,…

## Slide 9
Effective gain vs mesh voltage is measured by using the fitted MPV of the (Landau) distribution of the measured charge. It compares well with direct gain measurements with a 55Fe source on a detector without resistive foil (x2 lower)

## Slide 10: Crack scan
Study the hit reconstruction efficiency at or near the crack: take data with beam, moving by steps of 2mm
Beam direction
Outside crack
In crack
Upper crack
Lower crack
Padrow number (3x24=72)
Hits with ADC>150, 1-track evts

## Slide 11: Crack scan
Number of hits vs x
Effect of cracks is felt in an area ~1cm around.
In the worst case, 50% of the hits are lost

## Slide 12: Brute force alignment
ROW NUMBER : 3x24 = 72
Displacement in mm

## Slide 15: Alignment
Fitted displacement :
Rotation :
df = -1.7 mrad, 3.7 mrad and 8.4 mrad
Translation:
dx = -54 µm, 180 µm and 278 µm
After alignment : only distortions remain (probably ExB effect)
Maximum effect of distortions 400µm
Z=10cm

## Slide 16: On-going electronics developments
GdSP : gaseous detector signal processing (or Go digital as Soon as Possible)
CERN-based collaboration lead by Paul Aspell, including Saclay: F. Guilloux and E. Delagnes, continuing S-ALTRO evolution.
Common development for CMS GEM µ-chambers and LC TPC. Recently considering ALICE upgrade (GEM TPC and µ chambers).
Main developments : 130 nm technology, 64 or 128 channels, low noise and ultra-low consumption, many power domains to ease power switching
First Si test to be submitted in February 2013 (AIDA)
Goals for the FE: for 10 pF detector capacitance and 100 ns peaking time ENC<900 e- and power < 1mW/ch

## Slide 17: Front End specifications
Parameter | VFAT2 (IBM 0.25) | SALTRO (IBM 0.13) | VFAT3 / GdSP (IBM 0.13)
Linear range | +- 12fC | 150fC | Max 200fC
Input capacitance (pF) | 20 | 0-20 | 5 – 10 – 30 – 60
Noise | ~500e-  40-60e-/pF @ 25ns | ~ 650e-  15e-/pF @ 120ns | < SAltro /VFAT
From Paul Aspell
Parameter | VFAT2 | SALTRO | VFAT3 / GdSP
Power (mW/channel) | 1.5 (IBM 250nm) (incl. comparator) | 10 | << 1
Starting from data measured on ABCN (130nm design by Jan Kaplon for ATLAS SCT):
* 800 e- @ 5pF, tp=22ns, 100µW
* assuming serie noise only + strong inversion for the input transistor
 we can hope 530 e- @10pF, tp= 100ns, 200µW
 Ultra Low power low noise design seems feasible (cf slide 6)

## Slide 18: Proposal for CFE
Groups involved
Slow Control
(SPI)
Registers
DAC
CSA 2
Calib.
PZ 2
CSA 1/2
PZ 1
CSA 1/2/3
Calib.
PZ 1/2
x6
x2
MUX 91
Sallen Key
DDF
MUX 31
Bias
Regulator
Shaper
Output
Buffer
Bypass
Bypass
DDF
CSA 2
Calib.
PZ 2
DDF
x8
CEA-IRFU
INFN-BARI
CERN

## Slide 19: Summary
Successful test of 6 modules at a time
Most integration questions addressed
Improvements needed in contacts of the flat connectors
Multi-modules aspects addressed: alignment, reconstruction
Next step: 7 modules high quality test, full calibration on the test bench, detailed analysis in 2013-2014
R&D projects: thinner meshes, other charge dispersion devices (ceramic with ruthenium oxide).
Then a larger module with smaller pads for the inner wheel.
Thanks to our DESY colleagues, especially Ralf Diener, Christophe Rosemann and Felix Müller