500kV DC gun for ILC
Hiroshima University/KEK
Masao KURIKI
LCWS2013, 11-15, November, 2013

---
Contents
- Introduction
- Photo-cathode 500kV DC gun
- 500kV DC gun for ILC
- Summary

---
Introduction

---
ILC Design parameter

---
ILC electron source (1)
NEA GaAs/GaAsP superlattice cathode.
200kV inverted type DC gun.
>80% pol. , 3.2nC/bunch at IP.

---
ILC electron source (2)
- NEA GaAs/GaAsP for polarization is not compatible to RF Gun; DC photo-cathode gun is the only solution for LC.
- Current density from 200kV gun is limited by space charge resulting ~1ns bunch length.
  - 1ns → 200ps : two SHBs (316.7 and 433.3MHz).
  - 200ps → 20ps : TW 1.3GHz buncher (b=0.75).
  - Solenoid field for focusing.
- It is desirable to increase HV for shorter initial bunch length.
  - Simplfy the bunching section.
  - Potentially less beam loss and less margin.
- Improving 200kV → 500kV, the bunch length becomes 4 times shorter.

---
500kV DC gun

---
HV DC gun program in Japan
- Photo-cathode HV DC gun R&D program has been carried out by a collaboration among JAEA, KEK, Nagoya U. and Hiroshima U. since 2007.
- This is a common effort for multiple projects: ERL light source, Laser Compton Scattering photon source(Quantum Beam project), and Linear Colliders.
- 500kV and high beam current operation has been confirmed in a test beam line.
- The gun is now under commissioning as the injector of Compact ERL at KEK which is a demo-machine for GeV class ERL light source.

---
Impact of HV on emittance
Beam emittance is grown by non-linear space charge.
High field of the gun makes the beam emittace better by preventing the emittace growth.
800pC
80pC
I. V. Bazarov and C. K. Sinclair,
PRSTAB 8, 034202 (2005) .

---
ERL Project in Japan
- The next generation SR light source based on 3 GeV ERL (PERL) has been proposed.
- PERL is extendable to X-ray FEL oscillator.
- For technical demostration, cERL (compact ERL) is now under commissioning.
- cERL: Laser Compton X/γ-ray
  sources.
Parameters	Value
Energy	35/245 MeV
Current	10mA
Emittance (norm.)	0.1-1.0 mm.mrad

---
500kV DC gun
- 500kV DC biased gun with ceramic insulator.
- High voltage and high field are essential for high brighness and low emittance beam generation.
- Guard rings on the ceramic insulator prevent HV breakdown and ceramic punch through.
- The guard rings are also effective to prenvent cocentration of the bias voltage in a small region.
500kV bias
guard
rings
ceramic
insulator
support
rod
cathode
mount

---
Segmented Struture with guard rings (1)
Segmented
- HV of DC gun has been limited by ceramic insulator destruction.
- The reason is considered to be field emission from the support rod and field concentration by discharge.
- To prevent these phenomena, we employed segmented structure with guard rings.
One insulator
field emitted
electrons
field emitted
electrons

---
Segmented Structure with guard rings (2)
- Guard rings shade the ceramic from electrons emitted from the support rod by field emission.
- The bias voltage is distributed equally among each segments.
- Charge up and voltage concentration resulting ceramic destruction is threfore strongly prevented.
One insulator
Segmented

---
HV and Beam test

---
The 500kV DC gun system
- A 500kV DC photo-cathode gun with a cathode preparation system.
- High pumping ability by IP and NEG for extremely low vacuum pressure, <1e-9 Pa.
- Surface field <10MV/m for less field emission.
|E| (MV/m)
0
10
2
300
400
Z(mm)
500
600
4
6
8

---
JAEA DC gun beam line
- HV and beam test was carried out at JAEADC gun beam line.
- Laser: 5W 532nm.
- sx=0.1mm.
- Water cooled beam
  dump which is cap-
  able to measure the
  beam current.
- Differential pumping
  to prevent cotami-
  nation to the cathode.
steering magnet
500-kV gun
lightbox
beam dump
radiation shield
bending magnet
differential pump
beam profile monitor
beam expander
solenoid

---
500keV Beam Operation
- After HV conditioning up to 550kV, 500kV was quickly achieved.
- > 2.0mA current was achieved with 1.5 W laser and 0.28% QE.
- Vacuum pressure at gun was maintained at UHV during the operation.
0
100
300
500
600
high voltage (kV)
current (mA)
0.0
0.5
1.0
1.5
2.0
2.5
200
400
HV
I
time (minutes)
0
2
6
1e-9
vacuum pressure (Pa)
1e-8
1e-7
1e-6
1e-5
1e-4
1
3
4
5
@gun
@beam dump

---
10mA Beam operation
- 10 mA beam operation has been demonstrated.
- The test was carried out at lower voltage (180keV) due to field emisson from dust which was fixed by wiping.
- 10mA was achieved , but a significant vacuum pressure escalation was observed.
Laser power = 1.6W
QE = 1.5%
HV
0
50
100
150
200
HV (kV)
I (mA)
0
2
4
6
8
10
12
current
time (min.)
0
5
10
15
20
1e-9
Pressure (Pa)
1e-8
1e-7
1e-6
1e-5
0.2e+5
0.4e+5
0.6e+5
0.8e+5
1.0e+5
0.0
radiation
vacuum@gun
vacuum@dump

---
Operational Lifetime
- Operational lifetime was examined by contineous beam operation.
- Beam current decay was observed showing QE degradation.
- Charge lifetime was extracted to be 48C.
- It is enough for LCs, but not sufficient for ERL light source.
Laser power = 0.37W
QE = 0.68%
0.0
0.2
0.4
0.6
0.8
1.0
1.2
current (mA)
HV
I
0
100
300
500
600
high voltage (kV)
200
400
0 10 20 30 40 50 60
time (minutes)
@gun
@beam dump
1e-9
vacuum pressure (Pa)
1e-8
1e-7
1e-6
1e-5
1e-4

---
Operation at cERL
- The developed gun has been moved to cERL at KEK.
- CERL is now focused on commissioning of the injector.

---
HV stability
- 10 of 12 segments of the insulator ceramic are used due to soot on two segments.
- The gun is operated in 390kV. Voltage per segment is almost equivalent to that of 500kV on 12 segments.
- HV is quite stable during operation.
- The vacuum pressure is also stable at 1.3e-9 Pa.

---
QE in operation
- NEA GaAs is used for the commissioning.
- The 1/e lifetime is 7 month.
- It is enough for the low current commissioning.

---
500kV DC gun for ILC

---
500kV DC Gun for ILC
- Space charge limited current is 25.9A (8.2A/cm2).
  - 500kV Bias voltage,
  - 0.1m cathode-anode gap,
  - 10mm radius of cathode,
- Bunch length for 4.8nC (50% margin) charge is 190ps instead of 1ns.
  - Bunching factor is only 10 for 20ps for L-band booster.
  - 1.3GHz buncher is enough for bunching according to TDR design.

---
Another limit: SCL
- Another limitation becomes dominant in this current density.
- SCL : Surface charge limit by photo-voltage effect.
- A part of electrons are captured by the surface potential and its raise the potential wall to vacuum.
K. Togawa, NIMA (414), pp431–445(1998)
G.A. Mulhollan, Phy. Lett. A 282 (2001)

---
Surface Charge Limit
(a) sample 1b(Na=0.5),
(b) sample 2a(Na=1.0),
(c) sample 3(Na=2.0).
- The photo-voltage effect is compensated by facilitating recombination between the electrons and holes.
- Heavy p-dopesd surface relaxes SCL.
- 5.0A/cm2 is achieved by 2.0x1019 doping density.
- It is sufficient for 120 or 200kV DC gun, but not sufficient for 500kV ILC gun which requires 9.0A/cm2.
K. Togawa, NIMA (414), pp431–445(1998)
G.A. Mulhollan, Phy. Lett. A 282 (2001)
InGaAsÐAlGaAs SLwith4x1019cm-3

---
Summary
- 500kV DC photo-cathode gun is developed for high birghtness
  electron beam applications.
- Stable operation at 500kV has been confirmed.
- 10mA beam generation has been demonstrated.
- The gun in cERL commissioning shows good performances.
- The technology is applicable for ILC; It omits potentially SHBs.
- Surface charge limit becomes siginificant again. We have to solve it.

---
500kV DC gun R&D group
N. Nishimoria), R. Nagaia), S. Matsubaa), R. Hajimaa), M. Yamamotob), Y. Hondab), T. Miyajimab), H. Iijimac), M. Kurikic), M. Kuwaharad)
a) JAEA, b) KEK, c) Hiroshima U., d) Nagoya U.