## Slide 1: Review of ATF Shift - 100413
1/13
Mikey Davis
11 April 2013

## Slide 2: Introduction
2/13
Mikey Davis
11 April 2013
The shift focused on the IP BPMs
We tested the reference cavity signal amplitude
Checked out all the 2nd stage down mixers
We then moved on to trying to understand the resolution
We have found two “types” of second stage
I don’t think we managed to get a resolution better than about 5 μm

## Slide 3: A Quick Upstream Notes
3/13
Mikey Davis
11 April 2013
I re tested the phase sensitivity of the BPMs
The phase sensitivity didn’t seem to have changed much
  P2 has a phase sensitivity of 0.11 μm per °
  P3 has a phase sensitivity of 0.02 μm per °
In March
  P2 had a phase sensitivity of 0.08 μm per °
  P3 had a phase sensitivity of 0.09 μm per °
This stability is quite positive

## Slide 4: Reference Cavity
4/13
Mikey Davis
11 April 2013
10 dB of attenuation has been removed from the ref cavity input to the 1st stage (leaving 20 dB)
We tested the reference cavity signal amplitude at a number of different charge levels
The amplitude remained constant, if under the suggested 1 dB output level

## Slide 5: Ref Cav Levels
5/13
Mikey Davis
11 April 2013
Remembering the suggested output level is 1 dBm
N.B. We worked with charge of 0.05 and an extra 10 dB of atten. in March
Laser Int (%) | ICT Charge (IP) | Limiter Out (p2p V) | Limiter Out (dBm)
20 | 0.21 | 400 | -4
15 | 0.21 | 400 | -4
10 | 0.23 | 400 | -4
5 | 0.33 | 400 | -4
4 | 0.29 | 400 | -4
3 | 0.25 | 400 | -4
2 | 0.2 | 400 | -4
1 | 0.16 | 400 | -4

## Slide 6: Ref Cavity
6/13
Mikey Davis
11 April 2013

## Slide 7: 2nd Stages
7/13
Mikey Davis
11 April 2013
We spent some time testing all the second stages
We performed a calibration over a relatively wide range at 30dB (to get large signal to noise) for every 2nd stage
N.B. For the 1st and 2nd attempt the inputs were switched
The consts. are un charge normalised
2nd Stage | 07-3 | 07-1 | 07-2 | 06-2 | 07-4 | 06-1
1st Attempt | 0.265 | -0.779 | 0.226 | -0.842 | -0.229 | 0.750
2nd Attempt | 0.243 | 1.09 |  |  |  | 

## Slide 8: 2nd Stages
8/13
Mikey Davis
11 April 2013
There are clearly two “types” of 2nd stage
We started off by selecting 06 – 1 and 06 – 2
  Young Im should explain this
We then proceeded to try and test the resolution with this
2nd Stage | 07-3 | 07-1 | 07-2 | 06-2 | 07-4 | 06-1
1st Attempt | 0.265 | -0.779 | 0.226 | -0.842 | -0.229 | 0.750
2nd Attempt | 0.243 | 1.09 |  |  |  | 

## Slide 9: Calibration Constants
9/13
Mikey Davis
11 April 2013
We undertook a series of measurements to try and test the resolutions
First calibrated at atten. of 30, 20, 10 and 0 dB
Then took a jitter run at each of these settings
Found that the jitter was still considerably higher in Honda than Epics
I have checked my answers against Young Im
  We get similar (at the 20pc level) calibration consts.

## Slide 10: Calibration Constants
10/13
Mikey Davis
11 April 2013

## Slide 11: Resolution Tests
11/13
Mikey Davis
11 April 2013
Resolution runs were then undertaken at each attenuation settings
Atten(dB) | EPICS | 30 | 20 | 10 | 0
IPA (μm) | 0.91 | 12.4 | 6.83 | 5.30 | 5.00
IPB (μm) | - | 14.3 | 6.59 | 5.28 | 4.94
Mean Pos | -0.9 | -69 | -27 | -3 | 0.6
Res (μm) | - | 13.0 | 5.88 | 0.52 | 0.46
EPICS Pos | - | 61 | 62 | 62 | 63

## Slide 12: Jitter Run
12/13
Mikey Davis
11 April 2013

## Slide 13: Resolution Tests
13/13
Mikey Davis
11 April 2013
We then swapped the second stages and used two with a higher sensitivity
Atten(dB) | 20 | 10
IPA (μm) | 5.39 | 4.72
IPB (μm) | 5.71 | 4.67
Mean Pos | 32 | 39
Res (μm) | 1.43 | 0.87
EPICS Pos | 108 | 109
Atten(dB) | EPICS | 30 | 20 | 10 | 0
IPA (μm) | 0.91 | 12.4 | 6.83 | 5.30 | 5.00
IPB (μm) | - | 14.3 | 6.59 | 5.28 | 4.94
Mean Pos | -0.9 | -69 | -27 | -3 | 0.6
Res (μm) | - | 13.0 | 5.88 | 0.52 | 0.46
EPICS Pos | - | 61 | 62 | 62 | 63