## Slide 1: Laser calibration systems
STAR, ALICE, NIKA, ILC?
A.Lebedev, BNL
LCTPC meeting, May 31, 2013

## Slide 2: Why we need laser?
Simulation of charged particles with infinite momentum in gas.
Drift velocity measurement.
Misalignment and global ExB effect correction.
Space charge buildup corrections.

## Slide 3: STAR TPC laser system conception

## Slide 4
. Optical scheme for one half TPC.
-Mirror bundles glued to pyrex tube,
serving as a light optical bench.
-
Front view of laser raft
Outer field cage

## Slide 5
Laser tracks in TPC (ideal)
Two sets of angles in mirror bundles-
wide and narrow.

## Slide 6
This is a picture with ~ 400 laser beams in STAR TPC.
Wide laser beam ~ 30mm diameter used to illuminate 1mm mirrors.

## Slide 7: Main results
High accuracy in drift velocity measurements – reached ~ 0.02%.
This accuracy in Z ~200-300 um allows to use TPC tracks to align and calibrate space position of silicon drift chambers. Usually silicon detectors used to calibrate TPC.
ExB effect correction.
SHORTCOMINGS
-we tried to produce laser beams crossed in the TPC space; very difficult and created problems in track reconstruction,
-lasers not used for sector’s alignment,
-wide lasers are useless in any analysis

## Slide 8
Drift velocity measurements, run 2012
Gas system failures due to thunderstorms

## Slide 9
Central membrane with Al stripes pattern
Membrane made from conductive capton
Al coated mylar strips were glued radially.
Membrane illuminated by diffuse UV-
by- product after passing laser raft.
Good contrast with photoefficiency-
Al/capton ~100.
Strips covers all sectors, we started to use
them for sectors alignment.
Work is in progress, accuracy could be
~ 50-100um.

## Slide 10: ALICE TPC laser system
In ALICE TPC laser system a STAR design widely was used,
BNL and MEPhI(Moscow) participated in optics and mechanical parts production –bundles mirrors holders,
Assembly, start-up and operation- NBI.
Main steps in laser construction-see
http://www.nbi.dk/~borge/tpclaser
Nuclear Instruments and Methods in Physics Research Section A, Volume 622, Issue 1, p. 316-367.

## Slide 11: ALICE TPC design pattern
Laser beams pattern in TPC.
Mirror bundle with narrow angle set

## Slide 12: Расположение оптики в ALICE TPC

## Slide 13
Position of bundles in ALICE TPC.
Bundles are shifted in Z to avoid crossings in TPC volume.
Design was accommodated to field cage structure.

## Slide 14
Side view of laser event in ALICE TPC
Lasers used to measure drift
velocity with accuracy ~150um
for drift distance ~ 250cm.

## Slide 15: Lasers for LC TPC
The same approach with mirror bundle.
Less beams for drift velocity measurement.
Simple wide laser beam fanout on TPC wheel at Outer Field Cage- NIKA TPC (DUBNA).
Separate laser system with thin Poisson beams
to align GEM(MICROMEGAS).

## Slide 16: NIKA TPC proposed laser beams
TPC electronics
Wide laser beam splitters

## Slide 17: Conclusions
Lasers are useful tool.
Detailed design is needed to accommodate previous experience to LC TPC.
Extensive R&D is necessary to implement Poisson beams into LC TPC.
A new method utilizing diffraction of UV laser beams on annular diaphragms provides very narrow laser beams with full diameter 100-400 mm, divergence ~0.05mrad and effective length up to 10 meters
(A. Lebedev, IEEE2007)