## Slide 1: Recent Micromegas TPC R&D results
David Attié, Paul Colas, Wenxin Wang, CEA/Irfu
+ contributions from Carleton, KEK, Kolkata
LCTPC Collaboration meeting
DESY, May 31, 2013

## Slide 2
Micromegas module
Module size: 23 cm × 17 cm
24 rows × 72 columns
Readout: 1726 Pads
Pad size: ~3 mm × 7 mm

## Slide 3
Resistive ink
~3 MΩ/□
Resistive Kapton
~5 MΩ/□
Standard
Resistive Kapton
~3 MΩ/□
Micromegas modules

## Slide 4
Z=5cm
Z=35cm
Z=50cm
Mean Residual vs Row Number
Z-independent distortions
Distortions up to 50 microns for resistive ink (blue points)
Rms 7 microns for CLK film
(red points)
Uniformity
Row Number

## Slide 5
Beam test
2008~2011:
Test with one module
T2K electronics
(32cm × 25cm × 14cm )
2012~2013:
Test with
multi-modules
Integrated electronics ( 23cm × 17cm × 5cm )
Fully integrated Module
Resistive Micromegas Detector
FEC
FEM

## Slide 6
Event display
2D or 3D display software available
Online monitor
Offline monitor

## Slide 7
Drift Velocity
Test in 2009 (B=0T):
T2K gas: Ar:CF4:iso=95:3:2
T=19oC - P = 1035hPa
H2O:35ppm
At E=230 V/cm :
Vdrift = 7.698 +- 0.040 cm/µs (Magboltz : 7.583+-0.025(gas comp.))
The difference is 1.5+-0.6 %

## Slide 8: Gain measurements with a small detector (bulk; T2K gas) and a calibrated chain
(at Saclay with Purba and Sudeb Bhattacharya, Supratik Mukhopadhyay and Najana Majumber)
Note that there is no resisitve foil and that the mesh signal is used

## Slide 9
Gain Study (with beam test data)
For 6.84 mm pitch and 5 GeV e-
Expectations for Ntot: 61.9
(from HEED simulation)
ADC
Heed
Data
Number of electrons
ADC

## Slide 10
Test bench
Energy Resolution :~13% rms
Gain ~2600
(using collimated source, Vmesh=380V)
The 55Fe source is put in an aluminum collimation tube and fixed on a set of two mechanical arms (X-Y).

## Slide 11
Test bench – Energy resolution
Energy Resolution :11~13% rms at 380V
(Gain ~2600)
(using collimated source)

## Slide 12
Spatial resolution
Event selection
Select events with exactly 1 track

## Slide 13
Spatial resolution
: supposed residual of the detector
: the fitted hit position in test row
: the position of track crossing the test row
X (mm)
X (mm)
(a) Fit track with all rows
(b) Fit track without using hit in row 35
(c) Fit track after selection hits (row 35)
Hits selection
Row number
Row number

## Slide 15
Spatial resolution
B=1 T Cd = 94.2 µm/√cm (Magboltz)

## Slide 16
Compare with GEM
( Spatial resolution )

## Slide 17
Field distortions
B=1T
B=0T
After alignment
B=0T

## Slide 18
Field distortions
Distortions at different drift distances
in B=0T or B=1T
B=1T
B=0T

## Slide 19: Mitigating distortions?
Leave the corner insulated and floating.
Burry the grounding of the resistive kapton.
(Not so easy)

## Slide 20
Conclusion
A lot of experience has been gained in building and operating Micromegas TPC panels.
The characteristics of the Micromegas modules, such as the uniformity, energy resolution, stability have been studied in detail.
7 modules have been successfully tested with full integration of the electronics at the same time. 9 modules have been manufactured and characterized in a quasi-industrial process.
Thanks to the resistive technology, the measured resolution is about 60 microns at zero drift distance with 3 mm wide pads. This meets ILC needs.
Distortions are observed. Their detailed study will allow to correct them and improve the design to minimize them.