## Slide 1: FONT digitisation studies of IP BPMs
D. Bett, N. Blaskovic, P. Burrows, G. Christian, M. Davis, Y. I. Kim, C. Perry
John Adams Institute, Oxford University
N. Blaskovic

## Slide 2: Introduction
Introduction to IP BPMs and electronics
Signal digitisation: waveforms and FFT
Calibrations and phase shifter operation
Dynamic range and steering the beam
Setup modifications
Q vs. beam/BPM tilt scans
Charge normalisation considerations
N. Blaskovic

## Slide 3: Location of IP BPMs
N. Blaskovic
IPC
IPA
IPB
IP

## Slide 4: IP BPMs on movers
N. Blaskovic
IP
IPA & IPB
IPC
movers
based on figure from N. Terunuma

## Slide 5: IP BPM signal processing
N. Blaskovic
IPA
IPB
IPC
Ref
splitter
electronics
electronics
electronics
attenuator
variable attenuator
diode
FONT5 board (digitiser)
port 1
port 2
port 1
port 2
port 1
port 2
I
Q
I
Q
I
Q
IP BPM C-band signal
Reference C-band signal
Base-band signal
6.4 GHz (y) / 5.7 GHz (x)
based on S. Jang

## Slide 6: FONT5 board
9 ADCs (analogue-to-digital convertors)
Sampling at 357 MHz (2.8 ns)
13 bit: ± 4095 ADC counts for ± 0.5 V
Based on a Xilinx Vertex 5 FPGA
N. Blaskovic

## Slide 7
N. Blaskovic
20 dB
IPB(Y) I
IPB(Y) Q
IPC(Y) Q
IPB(X) I
IPC(X) I
IPC(X) Q
IPC(Y) I
Ref(Y)
IPB(X) Q

## Slide 8
N. Blaskovic
20 dB
IPB(Y) I
IPB(Y) Q
IPC(Y) Q
IPB(X) I
IPC(X) I
IPC(X) Q
IPC(Y) I
Ref(Y)
IPB(X) Q

## Slide 9: Calibration
IP BPM mover exercised across dynamic range
Dynamic range given as ± 3.6 um at 0 dB with charge of 5x109 for current electronics gain and single-port BPM (T. Tauchi)
BPM movers: ~30 um/V (O. Blanco)
N. Blaskovic

## Slide 10
N. Blaskovic
20 dB
9 mover steps with 25 pulses per step
I and Q charge normalised using reference cavity

## Slide 11
N. Blaskovic
20 dB
I’ is proportional to y position
Q’ is a measure of ‘unwanted’ signals and beam pitch y’ through the BPM
I’
Q’
θIQ

## Slide 12
N. Blaskovic
20 dB
I’
Q’
θIQ

## Slide 13: Phase shifter
The IQ plot can be rotated in hardware by using the phase shifters in the electronics
Procedure:
  Phase shifter setting changed
  Calibration θIQ determined at each setting
  Plotting θIQ vs. phase shifter setting allows θIQ to be set to zero (i.e. I’ = I and Q’ = Q)
N. Blaskovic

## Slide 14
N. Blaskovic
x

## Slide 15
N. Blaskovic
y

## Slide 16: Dynamic range
By inspecting the waveforms over a full dynamic range calibration, the I and Q dynamic range is < ± 1000 ADC counts
N. Blaskovic

## Slide 17
N. Blaskovic
IPC(Y) I
IPC(Y) Q
Ref(Y)

## Slide 18: Minimising signals
To avoid electronics saturation, signals were minimised as follows:
N. Blaskovic
Minimise | Method | 
 | using beam | using BPM mover
x position | Move QD0FF(x) | Use x-mover
y position | Move QD0FF(y) | Use y-mover
y’ pitch | Move QF7FF(y) | Change pitch

## Slide 19: Variable attenuator
The variable attenuator on the raw C-band signals from the IP BPMs was varied from 50 dB to 0 dB
IPC calibrated at each setting, with waist in y on IPC
Checked dependence on attenuation:
  Calibration constant (I’ per mover offset)
  Jitter (standard deviation of position)
N. Blaskovic

## Slide 20
N. Blaskovic
Factor 2 disagreement at 10 dB

## Slide 21
N. Blaskovic
True jitter?
Resolution?
Resolution at 0 dB?

## Slide 22
N. Blaskovic
IPC(Y) I
IPC(Y) Q
Ref(Y)
10 dB
large
~1000
large
~1000

## Slide 23: Changes to the set-up
Over the course of 5 shifts, performed the following changes cumulatively:
  Changed IPB (Y) and IPC (Y) from 1 to 2 port operation by using external 180º hybrids
  Placed waist in x on IPC (as well as in y)
  Interchanged IPB (Y) and IPC (Y) electronics
All changes undone at end of operation
N. Blaskovic

## Slide 24: IP BPM signal processing
N. Blaskovic
IPA
IPB
IPC
Ref
splitter
electronics
electronics
electronics
attenuator
variable attenuator
diode
FONT5 board (digitiser)
port 1
port 2
port 1
port 2
port 1
port 2
I
Q
I
Q
I
Q
IP BPM C-band signal
Reference C-band signal
Base-band signal
6.4 GHz (y) / 5.7 GHz (x)
hybrid
hybrid

## Slide 25: Waist scan
Waist scan in x at IPC
Performed by changing QF1FF current
N. Blaskovic

## Slide 26: Waist scan
Waist scan in y at IPC
Performed by changing QD0FF current
N. Blaskovic

## Slide 27: Variable attenuator
Variable attenuator varied from 50 dB to 0 dB with all changes implemented
N. Blaskovic

## Slide 28
N. Blaskovic
Agreement at 0 dB
Agreement at 10 dB

## Slide 29
N. Blaskovic
True jitter?
Resolution?
Resolution at 0 dB?

## Slide 30
N. Blaskovic
IPC(Y) I
IPC(Y) Q
Ref(Y)
10 dB
smaller
~300
smaller
~500

## Slide 31: Calibration ranges
Comparison of 3 calibrations at 0 dB over
  T. Tauchi’s dynamic range / 2
  T. Tauchi’s dynamic range
  T. Tauchi’s dynamic range x 3
N. Blaskovic

## Slide 32
N. Blaskovic
Tauchi/2

## Slide 33
N. Blaskovic
Tauchi
Saturation
or drift?
Saturation
or drift?

## Slide 34
N. Blaskovic
Tauchi*3

## Slide 35: Q vs. pitch scan
Dependence of Q on relative beam pitch (y’ ) to BPM axis tested by
  Changing beam pitch y’ using QF7FF mover
  Changing BPM pitch using IPC mover E:
N. Blaskovic
Mover D
Mover C
Mover E
beam
BPM block C (from above)
figure from O. Blanco

## Slide 36
N. Blaskovic
Operation on FONT IP BPM shifts

## Slide 37
N. Blaskovic
Operation on FONT IP BPM shifts

## Slide 38: Q vs. pitch scan method
N. Blaskovic
Move QF7FF(Y)
Centre beam using QD0FF(Y) mover
Perform calibration using mover

## Slide 39: Reference diode
Comparison of reference diode to sum (charge) signal of MFB1FF stripline BPM
The two charge indicators show a linear dependence, but are not proportional
May lead to incorrect charge normalisation
Also, reference signal is small
N. Blaskovic

## Slide 40
N. Blaskovic

## Slide 41: Further work
Investigate use of other signals for charge normalisation, e.g. stripline BPM or ICT
Use band pass filters to remove unwanted frequencies
N. Blaskovic

## Slide 42: Issues
Large ~40 MHz ripple at low attenuations
Large jitter measured even on waist
Apparent beam drift
Small reference signal
Reference signal not proportional to stripline charge measurement
N. Blaskovic

## Slide 43: Conclusions
IP BPM signals digitised by FONT5 board
Minimised I, Q signals by beam steering
Calibrated from 50 dB to 0 dB
Progressed to achieve calibration constants that scale with attenuation
Ripples, apparent beam drift and limited charge normalisation require attention
N. Blaskovic