## Slide 1: Multi-parameter fit
FINDING 3BPM RESOLUTION
Talitha

## Slide 2: Siwon’s method

## Slide 3: Siwon’s method

## Slide 4: Siwon’s method
IPBPM-A
IPBPM-B
IPBPM-C
My results using Siwon’s method of calculation:

## Slide 5: Multi-parameter fit
All I and Q values [counts] normalised by reference signals.
Example for predicting IPA-YI'
IPA-YI' = a1IPB-YI' + a2IPB-YQ' + a3IPB-XI + a4IPB-XQ + a5IPC-YI' + a6IPC-YQ' + a7IPC-XI + a8IPC-XQ + a9IPA-XI + a10IPA-XQ + a11Ref-Y + a12Ref-X + a13
ResolutionIPA-Y = std(IPA-YI'meas – IPA-YI'pred)
kIPA-Y × GF
k = calibration factor [counts][mm]-1
GF = geometric factor

## Slide 6: When to apply calibration factor?
Siwon/Honda: perform fit using I and Q signals in ADC counts. Apply calibration from ADC counts into nm at the very last stage after fit has been applied, using just the calibration factor for the BPM currently calculating resolution for, which is found from a calibration run of that BPM.
FONT: perform fit using calibrated signals converted to positions. Apply calibration from ADC counts into nm before fitting, converting all I’ signals into positions using the three different calibration factors for each BPM determined from separate calibration runs.

## Slide 7: Calibrating before and after fitting
jitRun28_0dB_030_ipbpm_150619 |  | FIT PARAMETERS |  |  | 
Fit description | IP BPM | (YI’)1 | ( YI’)2 | Constant | Resolution (nm)
2 parameters - apply calibration factor before performing fit: integration 53:63 | A | 1.377 | -0.432 |  | 173.6
 | B | 0.677 | 0.358 |  | 169.7
 | C | -1.355 | 2.283 |  | 190.1
2 parameters - apply calibration factory after performing fit: integration 53:63 | A | 1.447 | -1.084 |  | 147.6
 | B | 0.645 | 0.855 |  | 145.9
 | C | -0.54 | 0.956 |  | 363.1
3 parameters - apply calibration factor before performing fit : integration 53:63 | A | 0.76 | -0.869 | -6.834 | 181.1
 | B | 0.308 | -0.198 | -5.87 | 165.4
 | C | -0.495 | -0.278 | -11.322 | 181.6
3 parameters - apply calibration factor after performing fit: integration 53:63 | A | 0.799 | -2.18 | -7.534 | 109.4
 | B | 0.293 | -0.472 | -6.159 | 154
 | C | -0.197 | -0.116 | -4.975 | 203.6
Geometric parameters | 
1.4638 | -0.4638
0.6831 | 0.3169
-2.1559 | 3.1559
When fitting with no constant and calibrating before fitting (first data set), parameters can be compared to the geometric parameters (above in blue) derived from the known separation of the BPMs.
IPA-YI' = a1IPB-YI’ + a2IPC-YI’ + a3
Integration 53:63 Calibrations |  | 
IP BPM | Theta | k
A | 0.053 +/- 0.005 | 1.102 +/- 0.026
B | 0.262 +/- 0.006 | 1.049 +/- 0.013
C | 0.392 +/- 0.015 | 0.439 +/- -0.011

## Slide 8: Adding parameters
jitRun28_0dB_030_ipbpm_ 150619 |  | FIT PARAMETERS |  |  |  |  |  |  |  |  |  |  |  |  | 
Fit description | IP BPM | YI' | YQ' | XI | XQ | YI' | YQ' | XI | XQ | XI self | XQ self | Ref X | Ref Y | Constant | Resolution (nm)
3 parameters | A | 0.799 |  |  |  | -2.18 |  |  |  |  |  |  |  | -7.534 | 109.4
 | B | 0.293 |  |  |  | -0.472 |  |  |  |  |  |  |  | -6.159 | 154
 | C | -0.197 |  |  |  | -0.116 |  |  |  |  |  |  |  | -4.975 | 203.6
Add Q 5 parameters | A | 1.515 | -0.832 |  |  | -1.524 | -2.114 |  |  |  |  |  |  | 5.031 | 36.3
 | B | 0.538 | 0.475 |  |  | 0.555 | 1.321 |  |  |  |  |  |  | -1.695 | 40
 | C | -0.402 | -0.071 |  |  | 0.416 | -0.595 |  |  |  |  |  |  | 2.048 | 127
Add X 9 parameters | A | 1.504 | -0.724 | -0.411 | 0.835 | -1.541 | -2.377 | -0.089 | 0.505 |  |  |  |  | 1.565 | 27.7
 | B | 0.556 | 0.051 | -0.026 | -0.117 | 0.587 | 1.969 | 0.021 | -0.147 |  |  |  |  | -0.474 | 28.3
 | C | -0.382 | -0.105 | 0.02 | -0.018 | 0.357 | -0.527 | -0.289 | 0.094 |  |  |  |  | 1.227 | 112.7
Add self X 11 parameters | A | 1.539 | -0.589 | 0.022 | 0.822 | -1.472 | -2.478 | 0.111 | 0.241 | -0.054 | 0.105 |  |  | -0.027 | 26.7
 | B | 0.541 | 0.189 | 0.029 | -0.079 | 0.534 | 1.732 | -0.031 | -0.14 | -0.008 | -0.616 |  |  | 1.143 | 27.7
 | C | -0.379 | -0.099 | -0.032 | -0.023 | 0.357 | -0.539 | 0.006 | 0.145 | 0.12 | 0.024 |  |  | 0.798 | 109.9
Add Ref X 12 parameters | A | 1.696 | -0.745 | -0.066 | -0.389 | -1.236 | -2.341 | 0.059 | -0.05 | 0.241 | 0.241 | 0.003 |  | 12.136 | 18
 | B | 0.545 | 0.492 | -0.053 | -0.117 | 0.555 | 1.263 | -0.008 | 0.047 | 0.199 | 0.199 | -0.002 |  | -5.423 | 19.4
 | C | -0.539 | -0.457 | 0.02 | 0.037 | 0.749 | -0.587 | -0.062 | -0.21 | -0.064 | -0.064 | 0.002 |  | 6.978 | 77.1
Add Ref Y 13 parameters | A | 1.752 | -0.993 | 0.067 | -0.027 | -1.159 | -1.921 | 0.071 | -0.009 | 0.031 | 0.224 | -0.001 | 0.003 | 13.693 | 17.3
 | B | 0.539 | 0.536 | -0.028 | -0.108 | 0.528 | 1.162 | -0.02 | 0.038 | 0.02 | 0.046 | 0 | -0.001 | -5.136 | 18.7
 | C | -0.611 | -0.471 | -0.007 | 0.049 | 0.93 | -0.745 | 0.015 | -0.007 | 0.079 | -0.037 | 0 | 0.002 | 8.725 | 65.6
Geometric factor (as per Siwon's method) | A | 1.752 | -0.993 | 0.067 | -0.027 | -1.159 | -1.921 | 0.071 | -0.009 | 0.031 | 0.224 | -0.001 | 0.003 | 13.693 | 30.1
 | B | 0.539 | 0.536 | -0.028 | -0.108 | 0.528 | 1.162 | -0.02 | 0.038 | 0.02 | 0.046 | 0 | -0.001 | -5.136 | 26.8
 | C | -0.611 | -0.471 | -0.007 | 0.049 | 0.93 | -0.745 | 0.015 | -0.007 | 0.079 | -0.037 | 0 | 0.002 | 8.725 | 28.9
All calibrating using integration between 53 and 63, and calibrating after the fit. All except the final analysis use a geometric factor that contains all fit parameters (except the constant).
IPA-YI' = a1IPB-YI' + a2IPB-YQ' + a3IPB-XI + a4IPB-XQ + a5IPC-YI' + a6IPC-YQ' + a7IPC-XI + a8IPC-XQ + a9IPA-XI + a10IPA-XQ + a11Ref-Y + a12Ref-X + a13

## Slide 9: Fitting with I’ and Q’ vs with I and Q
Compare fitting:
IPA-YI' = a1IPB-YI' + a2IPB-YQ’ + a5IPC-YI' + a6IPC-YQ’ + a13
jitRun28_0dB_030_ipbpm_150619 |  | FIT PARAMETERS |  |  |  |  | 
Fit description | IP BPM | YI' | YQ' | YI' | YQ' | Constant | Resolution (nm)
Add Q - 5 parameters - apply calibration factor after performing fit: integration 53:63 | A | 1.515 | -0.832 | -1.524 | -2.114 | 5.031 | 36.3
 | B | 0.538 | 0.475 | 0.555 | 1.321 | -1.695 | 40
 | C | -0.402 | -0.071 | 0.416 | -0.595 | 2.048 | 127
jitRun28_0dB_030_ipbpm_150619 |  | FIT PARAMETERS |  |  |  |  | 
Fit description | IP BPM | YI | YQ | YI | YQ | Constant | Resolution (nm)
Using I and Q instead of I' and Q' | A | 1.678 | -0.41 | -0.602 | -2.536 | 5.031 | 36.3
 | B | 0.512 | 0.503 | 0.009 | 1.433 | -1.695 | 40
 | C | -0.398 | -0.093 | 0.556 | -0.467 | 2.048 | 127
IPA-YI' = a1IPB-YI + a2IPB-YQ + a5IPC-YI + a6IPC-YQ + a13
Two different methods result in identical resolutions, just different fit parameters.

## Slide 10: Resolution vs sample # - 13-multifit

## Slide 11: Single sample integration
When fitting with no constant and calibrating before fitting (first data set), parameters can be compared to the geometric parameters (above in blue) derived from the known separation of the BPMs.
IPA-YI' = a1IPB-YI’ + a2IPC-YI’ + a3
jitRun28_0dB_030_ipbpm_150619 |  | FIT PARAMETERS |  |  | 
Fit description | IP BPM | YI' | YI' | Constant | Resolution (nm)
2 parameters - apply calibration factor before performing fit: single sample 58 | A | 1.466 | -0.34 |  | 224.5
 | B | 0.591 | 0.308 |  | 214.4
 | C | -1.019 | 2.295 |  | 259.7
2 parameters - apply calibration factory after performing fit: single sample 58 | A | 1.593 | -0.778 |  | 199.3
 | B | 0.544 | 0.65 |  | 196.5
 | C | -0.445 | 1.088 |  | 454.8
3 parameters - apply calibration factor before performing fit : single sample 58 | A | 0.324 | -1.157 | -12.39 | 180.1
 | B | 0.22 | -0.247 | -6.185 | 220.1
 | C | -0.56 | -0.176 | -11.499 | 169
3 parameters - apply calibration factor after performing fit: single sample 58 | A | 0.353 | -2.653 | -2.231 | 98.6
 | B | 0.202 | -0.52 | -1.025 | 202.4
 | C | -0.244 | -0.084 | -0.903 | 189.7
Single Sample Number 58 Calibrations |  | 
IP BPM | Theta | k
A | 0.112 +/- 0.007 | 0.180 +/- 0.004
B | 0.380 +/- 0.008 | 0.166 +/- 0.002
C | 0.574 +/- 0.019 | 0.079 +/- -0.002
Geometric parameters | 
1.4638 | -0.4638
0.6831 | 0.3169
-2.1559 | 3.1559

## Slide 12: Adding parameters
All calibrating using single sample 58, and applying calibration after the fit. All except the final analysis use a geometric factor that contains all fit parameters (except the constant).
jitRun28_0dB_030_ipbpm_150619 |  | FIT PARAMETERS |  |  |  |  |  |  |  |  |  |  |  |  | 
Fit description | IP BPM | YI' | YQ' | XI | XQ | YI' | YQ' | XI | XQ | XI self | XQ self | Ref X | Ref Y | Constant | Resolution (nm)
3 parameters | A | 0.353 |  |  |  | -2.653 |  |  |  |  |  |  |  | -2.231 | 98.6
 | B | 0.202 |  |  |  | -0.52 |  |  |  |  |  |  |  | -1.025 | 202.4
 | C | -0.244 |  |  |  | -0.084 |  |  |  |  |  |  |  | -0.903 | 189.7
Add Q - 5 parameters | A | 1.646 | -0.591 |  |  | -0.591 | -0.591 |  |  |  |  |  |  | -0.591 | 23.8
 | B | 0.484 | -0.133 |  |  | -0.133 | -0.133 |  |  |  |  |  |  | -0.133 | 32.5
 | C | -0.382 | -0.625 |  |  | -0.625 | -0.625 |  |  |  |  |  |  | -0.625 | 147.1
Add X - 9 parameters | A | 1.518 | 0.27 | -0.117 | 0.112 | -1.425 | -3.993 | 0.185 | 0.032 |  |  |  |  | -0.757 | 23
 | B | 0.513 | 0.274 | 0.015 | 0.035 | 0.521 | 1.789 | -0.021 | -0.106 |  |  |  |  | -0.027 | 29.7
 | C | -0.366 | -0.308 | 0.017 | -0.013 | 0.263 | -0.187 | -0.235 | -0.003 |  |  |  |  | -0.555 | 132.4
Add self X - 11 parameters | A | 1.53 | 0.199 | -0.017 | 0.229 | -1.428 | -4.001 | 0.197 | 0.091 | -0.043 | -0.049 |  |  | -0.742 | 22.7
 | B | 0.509 | 0.319 | 0.057 | 0.046 | 0.507 | 1.706 | -0.06 | -0.125 | -0.113 | -0.229 |  |  | 0.116 | 29.7
 | C | -0.366 | -0.293 | -0.036 | -0.02 | 0.28 | -0.22 | 0.04 | 0.084 | 0.159 | 0.019 |  |  | -0.581 | 125.3
Add Ref X - 12 parameters | A | 1.893 | -0.825 | -0.124 | -0.112 | -1.171 | -2.998 | 0.085 | 0.038 | 0.019 | -0.003 | 0.001 |  | 2.051 | 20.7
 | B | 0.475 | 0.517 | 0 | 0.008 | 0.381 | 1.307 | -0.01 | -0.023 | 0.008 | 0.051 | 0 |  | -0.503 | 23.1
 | C | -0.467 | -0.383 | -0.02 | -0.009 | 0.597 | -0.527 | 0.008 | -0.011 | 0.128 | 0.012 | 0 |  | 0.334 | 98.6
Add Ref Y - 13 parameters | A | 2.03 | -1.278 | 0.006 | 0.11 | -1.077 | 0.074 | 0.074 | 0.093 | -0.027 | -0.002 | 0 | 0.001 | 2.949 | 19.4
 | B | 0.472 | 0.508 | 0.01 | 0.007 | 0.368 | -0.014 | -0.014 | -0.031 | -0.018 | -0.005 | 0 | 0 | -0.511 | 22.4
 | C | -0.512 | -0.332 | -0.034 | -0.007 | 0.75 | 0.048 | 0.048 | 0.068 | 0.124 | 0.036 | 0 | 0 | 0.859 | 85.6
Geometric factor (as per Siwon's method) | A | 2.03 | -1.278 | 0.006 | 0.11 | -1.077 | 0.074 | 0.074 | 0.093 | -0.027 | -0.002 | 0 | 0.001 | 2.949 | 39.6
 | B | 0.472 | 0.508 | 0.01 | 0.007 | 0.368 | -0.014 | -0.014 | -0.031 | -0.018 | -0.005 | 0 | 0 | -0.511 | 32.5
 | C | -0.512 | -0.332 | -0.034 | -0.007 | 0.75 | 0.048 | 0.048 | 0.068 | 0.124 | 0.036 | 0 | 0 | 0.859 | 34.6
IPA-YI' = a1IPB-YI' + a2IPB-YQ' + a3IPB-XI + a4IPB-XQ + a5IPC-YI' + a6IPC-YQ' + a7IPC-XI + a8IPC-XQ + a9IPA-XI + a10IPA-XQ + a11Ref-Y + a12Ref-X + a13

## Slide 13: December 714 MHz BPF run
714 MHZ BPF DECEMBER: resol_0dB_8_FB_off |  | FIT PARAMETERS |  |  |  |  |  | 
Resolution run | IP BPM | YI' | YQ' | YI' | YQ' | Ref Y | Constant | Resolution (nm)
3 parameters - Integration 53:63 - Calibration after fitting | A | -1.228 |  | 0.272 |  |  | 3.13 | 53.4
 | B | -0.741 |  | 0.195 |  |  | 2.237 | 56.6
 | C | 3.119 |  | 3.707 |  |  | -10.362 | 37.6
Add Q' 5 parameter fit - Integreation 53:63 - Calibration after fitting | A | -1.3 | -0.674 | 0.25 | -0.33 |  | 1.326 | 35.3
 | B | -0.731 | 0.29 | 0.189 | -0.187 |  | 2.001 | 40.2
 | C | 3.262 | -2.057 | 4.043 | 1.081 |  | -6.943 | 25.5
Add Ref Y - 6 paramter fit - Integration 53:63 - Calibration after fitting | A | -1.276 | -0.455 | 0.252 | -0.312 | 0.000135 | 2.334 | 35.8
 | B | -0.74 | 0.188 | 0.188 | -0.212 | 0.000068 | 2.305 | 40.1
 | C | 3.261 | -2.078 | 4.047 | 1.103 | 0.000022 | -6.801 | 25.4

## Slide 14: IPA calibration with C-band BPFs

## Slide 15: IPB calibration with C-band BPFs

## Slide 16: IPC calibration with C-band BPFs