## Slide 1: Phase Monitor Calibrations Vs. Sample No.
Jack Roberts

## Slide 2: Brief Phase Monitor Intro
Mon1
Mon2
Mon3
Measure arrival time of beam compared to ideal.
Expressed as phase: 360 degrees of 12 GHz ~80 pico seconds.
3 phase monitors: 2 upstream (Mon1 and Mon2), 1 downstream (Mon3).

## Slide 3: Rough Phase Monitor Electronics Schematic
Σ
Mixer ~ Asin(θ)
Diode~ A^2
LO
(12 GHz Reference phase)
Phase Shifter
Phase Monitor
e-
Hybrid
Phase Monitor Electronics (mixers)
Calibration procedure: Measure mixer output for a range of phase shifter values.

## Slide 4: Calibration Procedure: Data Taking
Phase shifters:
  1 x phase shifter, dial with readings in 1 GHz degrees (Mon2).
  2 x phase adjusters, no reading so count turns of nut (Mon1 and Mon3).
Usually take 20 pulses at these shifter settings (~10 points total):
  Mon1 and Mon3: 100 – 130o in steps of 3o(360o of 12 GHz in steps of 36o).
  Mon2: 0 to 16.75 turns of the nut in steps of 2 turns (~260o in steps of ~31o).

## Slide 5: Calibration Procedure: Data Processing
Convert raw signals in SiS counts to volts.
Subtract mean baseline noise.
Pick a sample range to use: usually take mean of 3 samples in a stable part of the pulse (same range for all monitors).
Remove any points >3 sigma away from mean at that shifter setting.
Calculate mean and std at each shifter setting.
  1/std^2 used as fit weights.

## Slide 6: Calibration Procedure: Fit (without diode)
Make initial estimate for parameters:
  A: (|maxMixer|+|minMixer|)/2
  b: (12/shifterFrequency)*(pi/180)
  c: based on mixer output at first point -> NOT WORKING CORRECTLY!
  d: (|maxMixer|-|minMixer|)/2
MatLab custom equation fit (uses finite differencing) with confidence intervals on parameters.
  Found that fit is failing in some cases this week, mostly for Mon3 and at least partly due to incorrect choice of initial value for c.
  Will fix / investigate difference methods.
Mixer: mixer output at each shifter setting.
A: amplitude of sine curve.
b: conversion between shifter setting and 12 GHz radians.
x: shifter setting
c: phase offset of sine at shifter=0.
d: zero offset of sine curve.
NB: Baseline noise already subtracted so d should be small.
A and d are most critical for phase reconstruction. I don’t normally use b and c after the fit.

## Slide 7: Initial Results (considering fit issues): 20151001
MON1

## Slide 8: Initial Results (considering fit issues): 20151001
MON2

## Slide 9: Initial Results (considering fit issues): 20151001
MON3
Pretty poor beam downstream.

## Slide 10: Fit Examples
MON1
MON2

## Slide 11: Fit Examples
MON3 - Good
MON3 - Bad

## Slide 12: Calibration Amplitude vs. Sample (A)

## Slide 13: Calibration Frequency vs. Sample (b)

## Slide 14: Calibration Phase Offset vs. Sample (c)

## Slide 15: Calibration Amplitude Offset vs. Sample (d)

## Slide 16: Summary
Need to work on the fitting procedure.
Initial impressions:
  Calibrations done with the phase adjusters are worse than with the “proper” shifter.
  Mon2 (with shifter) calibration vs. sample no. seems fairly flat along the pulse.