## Slide 1
N. Blaskovic, T. Bromwich, R. Ramjiawan
Dynamic Range and Resolution

## Slide 2: Average signal per mover setting
I signal averaged for each mover setting.
The smallest I signals correspond to mover settings close to 8.3V.
The mover setting that centralises the beam within the BPM is at the edge of the mover range for this file.
At samples 55 and 59 signals at different mover settings cross.
Sample number
I Signal
8.3V
4.0V
IPCyCal6_10dB_ipbpm_160617.mat

## Slide 3
4.0V
Larger I signals
8.3V
Smaller I signals
Edge of the Mover Range
I signal flattening at higher mover values.
Mover settings for these triggers are reaching the edge of range 0-8.4V.
IPCyCal6_10dB_ipbpm_160617.mat

## Slide 4: IPC mover
The linear range appears to be >2V which is twice as large as the expected dynamic range.
Expected dynamic range: 5m (10dB, charge, from Honda’s paper) 1.19V (10dB, charge.)
4.0V
8.3V
>2V

## Slide 5
Sample 50
Sample 51
Sample 52
Sample 53
Sample 54
Sample 55
Sample 56
Sample 57
Sample 58
Sample 59
Sample 60
Sample 61

## Slide 6: Saturation?
I signal turns non-linear at mover step 6 for both sample numbers despite the I signal being two very different values.
Not saturation for sample 51 as will not be saturating at 250 ADC counts.
Sample 51
Sample 52
300
0
2200
400

## Slide 7: Three Resolution Methods
‘Complete fitting method’:with residuals, and found from fitting.
‘Complete geometric method’: with residuals, and found using geometry.
‘Mixed fitting method’: with residuals found using fitting, and with and found using geometry.

## Slide 8
Resolution 10dB
Resolution method | Position | IP A | IP B | IP C
Fitting with GF from geometry |  | 0.126±0.008 | 0.146±0.008 | 0.094±0.005
Fitting with GF from geometry |  | 0.160±0.011 | 0.207±0.014 | 0.099±0.005
Geometric method |  | 0.192±0.051 | 0.192±0.051 | 0.192±0.051
Geometric method |  | 0.213±0.060 | 0.213±0.060 | 0.213±0.060
Fitting with fitted GF |  | 0.109±0.067 | 0.122±0.076 | 0.233±0.130
Fitting with fitted GF |  | 0.151±0.091 | 0.192±0.117 | 0.310±0.203
Resolution method | Position | IP A | IP B | IP C
Fitting with GF from geometry |  | 0.126±0.008 | 0.146±0.008 | 0.094±0.005
Fitting with GF from geometry |  | 0.160±0.011 | 0.207±0.014 | 0.099±0.005
Geometric method |  | 0.192±0.051 | 0.192±0.051 | 0.192±0.051
Geometric method |  | 0.213±0.060 | 0.213±0.060 | 0.213±0.060
Fitting with fitted GF |  | 0.109±0.067 | 0.122±0.076 | 0.233±0.130
Fitting with fitted GF |  | 0.151±0.091 | 0.192±0.117 | 0.310±0.203
160617_owl, samples 54-57, same files as last meeting.

## Slide 9
Resolution 20dB
Resolution method | Position | IP A | IP B | IP C
Fitting with GF from geometry |  | 0.804±0.045 | 0.579±0.31 | 0.249±0.013
Fitting with GF from geometry |  | 0.866±0.049 | 0.578±0.031 | 0.251±0.013
Geometric method |  | 1.043±0.030 | 1.043±0.030 | 1.043±0.030
Geometric method |  | 1.032±0.028 | 1.032±0.028 | 1.032±0.028
Fitting with fitted GF |  | 1.097±0.056 | 0.734±0.037 | 0.963±0.051
Fitting with fitted GF |  | 1.41±0.073 | 0.733±0.037 | 0.942±0.050
Resolution method | Position | IP A | IP B | IP C
Fitting with GF from geometry |  | 0.804±0.045 | 0.579±0.31 | 0.249±0.013
Fitting with GF from geometry |  | 0.866±0.049 | 0.578±0.031 | 0.251±0.013
Geometric method |  | 1.043±0.030 | 1.043±0.030 | 1.043±0.030
Geometric method |  | 1.032±0.028 | 1.032±0.028 | 1.032±0.028
Fitting with fitted GF |  | 1.097±0.056 | 0.734±0.037 | 0.963±0.051
Fitting with fitted GF |  | 1.41±0.073 | 0.733±0.037 | 0.942±0.050
160617_owl, samples 54-57, same files as last meeting.