## Slide 1: Diffraction Radiation present and future studies at ATF2
M. Bergamaschi1,2, A. Aryshev3, P. Karataev1, K. Kruchinin 1, R. Kieffer2, T. Lefevre2, S. Mazzoni2, N. Terunuma3
LCWS 2017, Strasbourg, 24th October 2017
John Adams Institute at Royal Holloway, Egham, United Kingdom
CERN European Organisation for Nuclear Research, Geneva, Switzerland
KEK High Energy Accelerator Research Organization, Tsukuba, Japan

## Slide 2
Outline
Project Overview
ODRI experiment at ATF2
Synchrotron Radiation Contribution at ATF2
Last operational year ODRI results at ATF2
Summary and outlook

## Slide 3
Motivations and Goals:
Develop a non-invasive transverse profile station for CLIC/ILC beams that can be scaled up
Develop, install and test a combined Optical Transition Radiation (OTR) and Optical Diffraction Radiation Interference (ODRI) emittance station at ATF2 at High Energy Accelerator Research Organisation (KEK)
To optimize sensitivity to smallest beam sizes, we plan to observe ODRI in the visible and far-UV wavelength range, down to approximately 200 nm
Overview

## Slide 4
Overview
Manpower |  | 
RHUL | P. Karataev | Scientific coordination Beam tests
 | M. Bergamaschi | Simulation, Beam test, Instrumentation
 | K. Kruchinin | Beam test, Instumentation
CERN | R. Kieffer | Beam test, Instrumentation
 | S. Mazzoni | Overall coordination Beam test
 | T. Lefevre | 
KEK | A. Aryshev | Beam test

## Slide 5
ODRI experiment location between QM14FF and QM13FF old laser-wire location
ODRI experiment at KEK ATF2

## Slide 6
Experiment installed at ATF2 in February 2016, in the laser-wire previous location where vertical beam can be focused to < 1um
ODRI experiment at KEK ATF2
Notes: Installation in Feb 2016

## Slide 7
ODRI experiment at KEK ATF2

## Slide 8
ODRI experiment at ATF2
The target as 4 slits for DR (50 to 201 µm)
A couple of vertical and horizontal mask slits can be inserted 13 cm upstream the target
ODR
Masks
TARGET
e- Beam
Synchronous Imaging and Angular acquisition for position filtering in angular

## Slide 9
Slit 1 201.7um
Slit 2 100.6um
Slit 3 77.2um
Slit 1 399um
Slit 2 198.9um
Slit 3 149.2um
TARGET
Mirrored areas
6x4mm2 around slits
MASK
Fully Mirrored
Slit 4 49.57um
Slit 4 100um
ODRI experiment at ATF2

## Slide 10
ODR in imaging
There is no visible beam size dependency of the pattern in imaging.
But the vertical position into the slit change the profile asymmetry => Optical Beam Position Monitor (BPM)
Offset

## Slide 11
ODR imaging as an optical BPM
Slit: 56um
Beam size: 1 µm
30 shots statistics
Position resolution: about 1 µm
Steering magnet to scan the beam inside the slit
Pmax
Pmin
Steering -0.85A
Asymmetry = (Pmin-Pmax)/(Pmax+Pmin)
Resolution ~1um
Range 45um

## Slide 12
ODR for beam size measurements
The beam size is extracted from the visibility Imin/Imax of the projected vertical component of the ODR angular distribution
An horizontal slit is used to measure a vertical beam size.
We use a polarizer to select only the vertically polarized ODR photons and 40nm BW filters to select the wavelength
The angular distribution is obtained using a camera located at the back focal plane of an optical lens (effective infinity)
ODR source
Aperture
Lens
CCD camera
On focal plane
d
f
Angular magnification
Polarizer
Filter

## Slide 13
ODRI at ATF2
Direct Image of the ODRI
2D Angular distribution of the ODRI

## Slide 14
Minimum too high = high SR Contribution
Synchrotron Radiation Contribution at ATF2
Observation of the angular pattern with target in the OTR position with mask inserted to evaluate Synchrotron Radiation Contribution

## Slide 15
OTR Angular pattern with steering magnet off and without mask
Strong Contribution, impossible to evaluate the TR angular pattern
Synchrotron Radiation Contribution at ATF2

## Slide 16
Although mask is inserted there is a strong Interference between Synchrotron Radiation and Transition Radiation because beam is not centered in the quadrupole (QM14FF) before the target.
Insertion of the mask (582 µm aperture)
Synchrotron Radiation Contribution at ATF2

## Slide 17
No mask present, Alignment of the quadrupole magnet present before the target (QM14FF)
Better profile but peak present in the middle of the pattern due to Synchrotron Radiation generated upstream in the beam-line
Synchrotron Radiation Contribution at ATF2

## Slide 18
Re insertion of the mask (582 µm)
Maximum reduction of the Synchrotron Radiation Contribution
Synchrotron Radiation Contribution at ATF2

## Slide 19
Before beam line optimization
After beam line optimization
Synchrotron Radiation Contribution at ATF2

## Slide 20
Before beam line optimization
After beam line optimization
Lower minimum = lower SR Contribution
Synchrotron Radiation Contribution at ATF2
Binning limit the minima => bigger angular magnification needed

## Slide 21
Synchrotron Radiation Contribution at ATF2
TR 2D profile without mask
TR 2D profile with mask

## Slide 22
ODRI 2D angular intensity
, beam size =

## Slide 23
[1]
With mask slit aperture, target slit aperture , and particle position at mask and slit, relativistic factor, observation wavelength, ratio between the particle velocity and the speed of light, , , , , , and angular coordinates in the observation plane.
ODRI simulation
[1] A. Cianchi et al., Non-intercepting electron beam size monitor using optical diffraction radiation interference, Phys. Rev. ST Accel. Beams 14,102803 (2011)

## Slide 24
, mask, target
Simulated Visibility vs Gaussian beam size for different target slit sizes
ODRI simulation
s
[2] P. Karataev, et al, Beam Size Measurements with Optical Diffraction Radiation at KEK Accelerator Test Facility, Phys. Rev. Lett. 93, 244802 (2004)
Projected Vertical Polarization
Component (PVPC) [2] of DR angular
Intensity for 2 beam sizes

## Slide 25
Quad current 120 (A)
Beam 42 um
Quad current 0 (A)
Beam 15 um
Quad current 40 (A)
Beam 24 um
Quad current 80 (A)
Beam 34 um
ODRI visibility Quad Scan
Filter:450nm Slit:201.7 µm Mask: 582 µm
PVPC of DR angular intensity for different beam sizes

## Slide 26
The side mirrors of the ODR slit were used to record reference OTR beam size measurement.
This could be done in CLIC/ILC using pilot beam (no target damage).
Then ODR can then be used for full beam charge.
OTR Beam size calibration

## Slide 27
ODRI angular intensity
,
PVPC of DR angular intensity
for 2 beam sizes
Measured Visibility for different beam sizes and mask/target combination

## Slide 28
ODR angular data vs simulation
,

## Slide 29
ODR angular data vs simulation
Synchrotron background Gaussian + 6 µm offset in the vertical alignment between slit and mask
,

## Slide 30
200nm Filter
450nm Filter
ODR angular data
Minimum beam size around 5 um (not calibrated, OTR PSF)
Last shifts results

## Slide 31
Summary
2017 ODRI Achievements
Imaging the DR source to monitor the position of the beam during angular data acquisition => beam in the slit centre
Sensitivity to beam size of tens of micrometers has been demonstrated, possible to descend to some microns
Beam orbit optimisation needed to minimize SR contribution
Mask contribution to block SR has been observed

## Slide 32
Outlook
Better measurement for smaller beam size
  Solution: bigger angular magnification new angular lens will be installed before November operation
Calibration with OTR PSF for micron size beam
  Solution: improvement of the present imaging optical line, bigger image magnification from 2 to 5, new optics will be installed before November operation
  November shifts -> commission new optics and acquire new data to calibrated sensitivity to small beam and comparison with last year measurement
  December shifts -> measurement in the visible and UV to confirm reproducibility of the system

## Slide 33
Thank you for your attention!