## Slide 1
Main RF power Coupler for European-XFEL
Walid KAABI- LAL/CNRS
Notes: How to edit the title slide
Upper area: Title of your talk, max. 2 rows of the defined size (55 pt)
Lower area (subtitle): Conference/meeting/workshop, location, date, your name and affiliation, max. 4 rows of the defined size (32 pt)
Change the partner logos or add others in the last row.

## Slide 2: Outlines:
Introduction
Fabrication process
RF conditioning at LAL
Troubles on the track

## Slide 3: Introduction:
Linac of 101 Cryomodules, equiped with 8 couplers each  Need of 808 Couplers.

## Slide 4: Introduction:
XFEL couplers are produced by 2 suppliers at 3 production sites:
Consortium Thales-RI (Thonon les bains-France and Koln-Germany): 670 units.
CPI (Beverly-Massachusetts-USA): 150 units.
Supported by DESY, LAL-Orsay has in charge:
The production monitoring and the quality control at Thales-RI sites.
The RF conditioning of all the couplers at Orsay and the weekly delivery of 8 couplers/week to IRFU-CEA (increased rate to 10 couplers/Week since January 2015) .

## Slide 5: Introduction:
Challenges:
First experience in industrial monitoring at a such big scale,
Successful know-how transfer to company: The clean room processes, TiN coating process & acceptance criteria,
Rescaling the RF process at LAL to ensure the conditioning of 8 couplers per week,
Respect of the overall project schedule.

## Slide 6
@ Thales-Thonon production site:
Step 1: Brazing of the inner and outer conductors:
Step 2: Copper plating of the inner surfaces:
Fabrication process:

## Slide 7
Inspection & control of all the produced parts at Thales-Thonon site.
Double check during the reception at RI site under the same inspection criteria.
Cold conductor
Warm Conductor
Fabrication process:

## Slide 8
Examples of observed copper defects:
 Qualitative & quantitative acceptance criteria definition: defects classified per types, dimensions et number. Agreement on corrective actions and validation tests.
Copper peel off
Blister
X 20
Pitting
X 30
X 22
Circular spots
Copper particle
Fabrication process:

## Slide 9
Copper thickness distribution in cold conductor (specif. 10μm ± 20% cylindrical parts & 10μm ± 30% in bellows relaxed to 10μm ± 50% ):
Fabrication process:

## Slide 10
RRR measuremets (Specif. 10<RRR<80)
Fabrication process:

## Slide 11
@ RI production site:
Fabrication process:

## Slide 12: Fabrication process:
RF Conditioning
Delivery to LAL

## Slide 13
RF conditioning at LAL:
WGBs assembly
Capacitors assembly
Actuators assembly

## Slide 14: RF conditioning at LAL:
RF Conditioning

## Slide 15: RF conditioning at LAL:
RF Conditioning

## Slide 16: RF conditioning at LAL:
RF Conditioning
Ligne D
Ligne C
Ligne B
Ligne A
PicD
PicC
PicA
PicB
PrcD
PrcC
PrcB
PrcA
PchC
PchB
PchA
V1A
V3A
V2A
V1B
V3B
V2B
V1C
V3C
V2C
V1D
V3D
V2D
PchD
Arcil1A
Arcil2A
Arcil1B
Arcil2B
Arcil1C
Arcil2C
Arcil2D
Arcil1D

## Slide 17: RF conditioning at LAL:

## Slide 18: RF conditioning at LAL:
Coupler pair disassembly & final inspection
Shipment to IRFU-CEA

## Slide 19: RF conditioning at LAL:
Summer Break 2014
Summer Break 2015
Xmas Break 2014
Xmas Break 2015
June 2016: The End!
CPI production ( 142 units):
- Average of 4 couplers/week
Thales-RI production (670 units):
- 8 couplers/week after ramp-up,
- Delivery increase to 10 couplers/week for 10 weeks (May 2015-Sept 2015)
24 Couplers delivered in 2013
-- Nominal delivery rate
-- Real rate

## Slide 20: RF conditioning at LAL:

## Slide 21: Troubles on the track: Dark layer on cold ceramic
Main defect discovered during the final inspection: Dark layer in cold ceramic window

## Slide 22: Troubles on the track: Light & arcing in the warm windows
Light intensity limit
Incident power signal
Reflected power signal
Light B1
Light C2

## Slide 23: Troubles on the track: Light & arcing in the warm windows

## Slide 24: Troubles on the track: Light & arcing in the warm windows
Light intensity limit

## Slide 25: Troubles on the track: Light & arcing in the warm windows
Module XM-1 couplers conditioning:
Light signal on spark sensors
Denis Kostin, MHF-sl, DESY

## Slide 26: Troubles on the track: Light & arcing in the warm windows
Light intensity limit
Sudden increase of the light intensity causing arc interlock: generally occurs at low power (under 250 kW) at the shortest pulse (20us)  impossible to go further: the pair is removed from the test stand.
Incident power signal
Reflected power signal

## Slide 27: Troubles on the track: Overheating during RF test
XM-1 Coupler 6 70K window overheating
(typical 70K window overheating case)
FPC simulation:
leak spot may have a contact-discharge
FPC push-rod leak:
leak spot discovered
FPC coax-gasket:
capacitors are replaced
E-XFEL FPC
All capacitors are replaced by coax gaskets: no push-rod leaks (RF is ON) since then
problem solved
CST MWS simulation
Courtesy to Denis Kostin

## Slide 28: Troubles on the track: Overheating during RF test
Denis Kostin, MHF-sl, DESY
XM-1 Coupler 6 70K window overheating
(typical 70K window overheating case)

## Slide 29: Troubles on the track: Overheating during RF test
Denis Kostin, MHF-sl, DESY
Coupler 70K window overheating due to loose connection between the WIC and the CP. Cold part must be grinded and cleaned. The WP must be changed.
Re-conditioning takes long time: cleaning is difficult, no baking, no US-wash...

## Slide 30: Troubles on the track: Overheating during RF test
WPs inspection after observed Overheating in module testes

## Slide 31: Troubles on the track: Overheating during RF test
Same “white traces” observed after conditioning at LAL, but the thermal sensor were not plugged to observe an overheating

## Slide 32: Summary:
850 power couplers were fabricated, RF conditioned and delivered to E-XFEL machine (808 assembled in modules and tested in AMTF + 42 spare units),
Coupler repair and parts replacement done in AMTF:
  35 warm parts replaced: inner screw contact problems, cold 70K window overheating and not “conditionable” discharges,
  30 push-rods replaced,
  Other repairs done (coupler movers, cables, WGs etc.);
All FPC capacitors are replaced by coax gaskets – no push-rod leaks anymore, problem well understood and fixed,
4 FPCs are shorted (not used) because of the coupler problems – cold 70K window overheating coupled with not conditionable discharge, two more couplers show T70K overheating effect – to be tested and decided.

## Slide 33
Janvier 2013
Septembre 2016