## Slide 1
ATF2 June Shifts
R. Ramjiawan
Friday, 29th June 2018

## Slide 2: Outline
First look at high-beta optics shift –
  Outline of shift
  Waveforms
  Calibrations with BPM movers (X & Y)
  Calibrations with AQD0FF mover (Y)
  Resolution measurements at nominal and large (but acceptable) position and tilt offsets

## Slide 3
Analysis from Shift 15/06/2018

## Slide 4: Shift Outline (15/06/18)
Owl and day shift: Resolution studies with X information, limiter phase information, bunch phase information from upstream board.
  Resolution scan: is resolution degraded my changing position/tilt within acceptable limits (i.e not in saturation)
  Resolution as function of charge
  Resolution as function of attenuation
  Resolution/limiter jitter as function of reference attenuation
High-beta optics
Beam jitter: IPA(X) 2.4 um, IPC(X) 2.6 um IPA(Y) 250 nm, IPB(Y) 180 nm, IPC(Y) 450 nm
QD0FF: 133.5A – waist between IPB and the IP.
Data taken at multiple different charges between ~0.3e10 and ~0.7e10.
This presentation: Scan of resolution within acceptable limits
  For data in this presentation: charge 0.6e10, baseline subtraction implemented on all I, Q channels.
  Samples analysed: 59:64, ref 67.
  Data sets have typically 400 triggers.
  Data synchronised using timestamps.

## Slide 5: Waveforms (X) - 20dB
IPA(X)
IPC(X)

## Slide 6: Waveforms (Y) – 10dB
IPA
IPB
IPC
The signal IPB(I) was highly insensitive to beam position and as such we weren’t able to reduce this throughout the shift.
Waveforms shown with baseline subtraction applied.

## Slide 7: IPA(X) Calibration
Samples analysed – 59:64, ref sample: 67.
IPxCal2

## Slide 8: IPC(X) Calibration
Samples analysed – 59:64, ref sample: 67.
IPxCal2

## Slide 9
Comparing BPM mover
And AQD0FF calibrations

## Slide 10: IPA(Y) Calibration (IPyCal9)
BPM mover
QD0FF mover
Samples 59:64, 67

## Slide 11: IPB(Y) Calibration (IPyCal9)
BPM mover
QD0FF mover
Samples 59:64, 67

## Slide 12: IPC(Y) Calibration (IPyCal9)
BPM mover
QD0FF mover
Samples 59:64, 67

## Slide 13
Limiter phase jitter
&
Bunch phase jitter

## Slide 14: Limiter Phase Jitter as Function of Sample
First sample: 59
Integrated window limiter jitter as function of sample window.
Line shows window used for analysis 59:64
Colin2_posNom_tiltNom1
Single sample limiter phase jitter

## Slide 15: Limiter phase jitter
Limiter phase jitter for the ‘resolution within acceptable limits’ scan and its repeat.
Charge decreased across course of scan.
Samples analysed 59:64, ref 67.
Average value: 0.034 radians.
23 % correlation between the limiter phase jitter and the mean bunch charge.
Consistent with previous measurements of 0.03 radians.

## Slide 16: Bunch Phase Jitter
Mean and standard deviation of bunch phase across the ‘resolution within acceptable limits’ scan and its repeat.
Sample used: 48 – corresponds to peak of Sum I waveform.
Error bars show bunch phase jitter.

## Slide 17: IPA Resolution-Fitting for calibration constant: k
Fitting to (Y) - nm
I'/q = 22.5 ± 1.1
I'/q and 1/q= 21.2 ± 1.1
I'/q, Q'/q and 1/q= 21.2 ± 1.1
I'/q, Q'/q, Self Q'/q and 1/q= 21.0 ± 1.1
I'/q, Q'/q, Self Q'/q, Lim. Phase and 1/q= 20.4 nm ± 1.0
I'/q, Q'/q, Self Q'/q, Bunch Phase and 1/q= 21.0 ± 1.1
Fitting to (X and Y) - nm
I'/q = 21.8 ± 1.1
I'/q and 1/q = 21.2 ± 1.1
I'/q, Q'/q and 1/q= 21.1 ± 1.1
I'/q, Q'/q, Self Q'/q, 1/q= 20.9 ± 1.1
I'/q, Q'/q, Self Q'/q, Lim. Phase, 1/q= 20.2 ± 1.0
I'/q, Q'/q, Self Q'/q, Bunch phase and 1/q= 20.9 ± 1.1
File: jitRun3
Ref sample: 67.
Geometric to Position Resolution= 23.2 ± 1.3 nm
Sample 55
Sample 65

## Slide 18: IPB Resolution-Fitting for calibration constant: k
Fitting to (Y) - nm
I'/q = 23.2 ± 1.0
I'/q and 1/q= 21.7 ± 1.0
I'/q, Q'/q and 1/q= 21.5 ± 1.0
I'/q, Q'/q, Self Q'/q and 1/q= 21.4 ± 1.0
I'/q, Q'/q, Self Q'/q, Lim. Phase and 1/q= 21.0 ± 1.0
I'/q, Q'/q, Self Q'/q, Bunch Phase and 1/q= 21.4 ± 1.0
Fitting to (X and Y) - nm
I'/q = 23.1 ± 1.0
I'/q and 1/q = 21.7 ± 1.0
I'/q, Q'/q and 1/q= 20.0 ± 0.9
I'/q, Q'/q, Self Q'/q, 1/q= 19.9 ± 0.9
I'/q, Q'/q, Self Q'/q, Lim. Phase, 1/q= 19.9 ± 0.9
I'/q, Q'/q, Self Q'/q, Bunch phase and 1/q= 19.9 ± 0.9
Geometric to Position Resolution= 23.2 ± 1.3 nm
File: jitRun3
Ref sample: 67.
Sample 55
Sample 65

## Slide 19: IPC Resolution-Fitting for calibration constant: k
Fitting to (Y) - nm
I'/q = 22.6 ± 1.5
I'/q and 1/q= 21.7 ± 1.5
I'/q, Q'/q and 1/q= 21.5 ± 1.5
I'/q, Q'/q, Self Q'/q and 1/q= 21.5 ± 1.5
I'/q, Q'/q, Self Q'/q, Lim. Phase and 1/q= 20.8 ± 1.4
I'/q, Q'/q, Self Q'/q, Bunch Phase and 1/q= 21.5 ± 1.5
Fitting to (X and Y) - nm
I'/q = 21.9 ± 1.5
I'/q and 1/q = 21.6 ± 1.5
I'/q, Q'/q and 1/q= 19.9 ± 1.4
I'/q, Q'/q, Self Q'/q, 1/q= 19.8 ± 1.3
I'/q, Q'/q, Self Q'/q, Lim. Phase, 1/q= 19.6 ± 1.3
I'/q, Q'/q, Self Q'/q, Bunch phase and 1/q= 19.8 ± 1.3
Geometric to Position Resolution= 23.2 ± 1.3 nm
File: jitRun3
Ref sample: 67.
Sample 55
Sample 65

## Slide 20
Scan of Acceptable
Positions and Angles

## Slide 21: Acceptable range of positions/angles
Scan through minimum, nominal and maximum acceptable positions.
Scan through minimum, nominal and maximum acceptable tilts.
Repeat.
Min position, max tilt | Nominal position, max tilt | Max position, max tilt
Min Position, nominal tilt | Nominal position, nominal tilt | Max position, nominal tilt
Min position, min tilt | Nominal position, min tilt | Max position, min tilt
I’/q across scan
Q’/q across scan

## Slide 22: Waveforms at min, nom, max of scan
Tilt Scan at nominal position
Position Scan at nominal tilt
Position and tilt scan
Min pos, min tilt
Min pos, max tilt
Max pos, min tilt
Max pos, max tilt
Min pos
Nom pos
Max pos
Min tilt
Nom tilt
Max tilt

## Slide 23: Scan 1 and Scan 2
Colour in heat-map corresponds to geometric resolution.
First scan: appear to be well centred about origin in both I’ and Q’
Second scan: no longer well centred and best resolution found at minimum tilt setting.
Scan 1
Scan 2

## Slide 24: Charge during the scan
Charge dropped across the course of the scan leading to a decreasing resolution (plots to the right).
79.1 % correlation
72.4 % correlation
Files from resolution scan that have geometric resolution of better than 60 nm

## Slide 25: Geometric resolution as function of pos. and tilt
Including data from scan 1 and scan 2.
Colour corresponds to geometric resolution.
Geometric resolutions are given as data labels in nm.
Res. Vs. Q’/q
Res. Vs. I’/q

## Slide 26: Improvement by using fitting
Q’/q
I’/q
Geometric resolution and resolution by fitting to just position shown below.
Plots on the right show the progression of Q’/q and I’/q across the scan.

## Slide 27: Beam position, angle drift across the scan
2.4 um
Geometric resolution across scan. Worst resolution settings not shown.
Nominal settings in red
Mean IPC I’/q across scan.
Nominal settings in red
Beam drift across the scan mostly seen in IPA and IPC not in IPB, possibly suggesting drift in angle rather than position.
Data shown here taken over ~2 hours.
IPB
IPA

## Slide 28: Improvement to resolution – fitting X
IPB resolution with fitting to:
IPA(X), IPA(Y), IPC(X), IPC(Y), charge compared with fitting to
IPA(Y), IPC(Y), charge.
Appears to have most impact at poorer resolution.

## Slide 29: Fitting for bunch phase, limiter phase
Showing the improvement to resolution gained from fitting out the bunch phase and limiter phase.
Fitting to Y data
Fitting to X and Y data

## Slide 30: Improvement from fitting lim. phase
Benefits from fitting limiter phase higher at settings with larger Q’/q value.
As expected from:
Plotted: Resolution fitting 7 parameters – resolution fitting 7 parameters & limiter phase.
(nm)

## Slide 31: Summary
BPM mover calibration agreed with AQD0FF calibration.
Resolution shows dependence on Q’/q within ‘acceptable limits of operation’.
Improvements by fitting to limiter phase seen for data sets with larger Q’/q.
  Limiter phase jitter ~0.034 radians.
No real improvement seen by fitting to the bunch phase.
  Bunch phase jitter ~0.0105 radians.
Slight improvement seen by fitting to X orbit for data sets with poorer resolution.

## Slide 32
Extra Slides

## Slide 33: Extra slide
ColinRun2 _posNomTiltNom1_10dB | 203, 203, 162, 26 nm (1/13)
ColinRun2 _posMinTiltNom1_10dB | 200, 200, 159, 27.7 nm (2/13)
ColinRun2 _posMaxTiltNom1_10dB | 206, 206, 165, 27.6 nm (3/13)
ColinRun2 _posNomTiltNom2_10dB | 203, 203, 162, 27.4 nm (4/13)
ColinRun2 _posNomTiltMin1_10dB | 173, 173, 280, 31.1 nm (5/13)
ColinRun2 _posMinTiltMin1_10dB | 170, 170, 277 (6/13)
ColinRun2 _posMaxTiltMin1_10dB | 176, 176, 283 (7/13)
ColinRun2 _posNomTiltMin2_10dB | 173, 173, 280, 29.4 nm (8/13)
ColinRun2 _posNomTiltMax1_10dB | 173, 173, 280 (9/13)
ColinRun2 _posNomTiltMax1_10dB | 233, 233, 44 (10/13)
ColinRun2 _posMinTiltMax1_10dB | 230, 230, 41 (11/13)
ColinRun2 _posMaxTiltMax1_10dB | 236, 236, 47 (12/13)
ColinRun2 _posNomTiltNom3_10dB | 203, 203, 162 (13/13)