## Slide 1: Analysis of SiW-ECAL technological prototype beam test with electron beam
Y. KatoB, K. GotoA, T. SueharaA
and ILD SiW-ECAL group
Kyushu UniversityA
The University of TokyoB
LCWS2019 @ Sendai, Japan
29th Oct. 2019
katou@icepp.s.u-tokyo.ac.jp
Notes: Hello everyone, this is Yu Kato, from the University of tokyo, and I will be talking about BSM Search using Higgs to Invisible Decay at the ILC.

## Slide 2: Table of Contents
Introduction
R&D of SiW-ECAL technological prototype
  FEV13-Jp Status
Beam Test 2019
Procedure for Energy Measurement
Analysis
  Trigger Adjustment
  Pedestal
  MIP
  TDC
  Shower
  Remain Issues

## Slide 3: International Large Detector
One of the detector concepts at the ILC
Optimized for Particle Flow Algorithm
  Reconstruct & identify all the particles
Components
  Vertex detector
  Trackers
  Calorimeters
    ECAL
      ScW-ECAL
      SiW-ECAL
    HCAL
  Muon Yoke
  etc.

## Slide 4: R&D of SiW-ECAL technological prototypes
TBD

## Slide 5: R&D of SiW-ECAL technological prototypes
TBD

## Slide 6: R&D of SiW-ECAL technological prototypes
Major changes in FEV11 → 13 and SMBv4 → v5
ASIC: SKIROC2 → 2A
  – Individual threshold control
  – Improvements on TDC
Smaller SMB footprint
Connection by 0.4mm-pitch flex cables
  – Two candidates, footprint compatible
FEV11 & SMBv4
FEV13 & SMBv5
Carbon
Carbon

## Slide 7: Analogue core: SKIROC2A
arXiv:1801.02024
Outputs: Gain (High/Low), TDC

## Slide 8: FEV13-Jp Status
ASIC: SKIROC2A
Si thickness: 320µm & 650µm New!
  256 ch/sensor × 4 sensor/slab
FEV-SMB Connection: Flexible cable or Micro-coaxial cable
EM shielding: w/ Carbon frame and cover
Power Pulsing
slab ID | P1 | P2 | P3 | K1 | K2
Si thickness | 650 µm | 650 µm | 320 µm | 650 µm | 650 µm
Assembly | in Kyushu U. | in Kyushu U. | in Kyushu U. | in LLR | in Kyushu U.
Total 5 slabs
in Kyushu U., Japan
See previous talk
for details

## Slide 9: Beam Test 2019 @ DESY
Beam time:
  24th June - 7th July at DESY test beam facility
  e- beam: 1 - 5 GeV
Presence from:
Support & Hardware from:

## Slide 10: Beam Test 2019 @ DESY
Beam time:
  24th June - 7th July at DESY test beam facility
  e- beam: 1 - 5 GeV
Objectives:
  Comparison of ASU based on BGA and based on Chip-On-Board (COB)
  Test of new SL-Boards
  Validation of FEV13-Jp ← Target of this talk
Programs:
MIP program (w/o Tungsten)
  Position scan for MIP calibration
  TDC test
  Angled beam: 25 deg.
  Retriggering / double pedestal
Shower program (w/ Tungsten)
  Energy measurement
  Response from large signal
  TDC / auto gain
  Edge effect

## Slide 11: Setup for Beam Test
Devices: 2 types of readouts
  DIF based slabs: FEV13-Jp × 5
  SLB based slabs:
    COB × 2
    FEV12 × 2
Absorber: Tungsten
  X0 = 3.5 mm, RM = 9 mm, λ0 = 96 mm
SLB slabs
DIF slabs
Tungsten (2.1 mm / 4.2 mm)
e-
e-

## Slide 12: Procedure for Energy Measurement
Single Slab Analysis
  Trigger adjustment & Masking of noisy channels
  Pedestal calibration
    16 chips × 64 channels × 15 memories
  Gain calibration using MIP
    16 chips × 64 channels
  (TDC calibration using test pulse)
    time walk correction
Multi Slab Analysis
  Timing coincidence
    using bunch crossing ID (BCID) Δt = 0.2 µs
  Event Building

## Slide 13: Trigger Adjustment
Threshold scan is performed for evaluation of S/NTrig and trigger adjustment.
Test pulse of {4.2, 8.4, 16.8} fC is injected.
  320 µm: 1 MIP ~ 4.2 fC
  650 µm: 1 MIP ~ 8.4 fC
S-curve is fitted by Err-function.
Preliminary
Injection [fC] | 4.2 | 8.4 |  | 16.8
mean [DAC] | 263.5 | 335.1 |  | 457.5
sigma [DAC] | 11.9 | 10.3 |  | 4.3
S/NTrig | 6.0 |  | 11.9 | 
Trigger is set as 0.5 MIP of 320 µm slab:
~ 230 DAC.

## Slide 14: Pedestal Analysis
Non-triggered ADC output (around ~300 [ADC])
Fitted by Gaussian
pedestal peak
(non-triggered)
MIP events
(triggered)

## Slide 15: Pedestal Homogeneity: Mean
Mean of Gaussian
SCA = 0 (Memory-cell dependence is referred later.)
Although there are differences between chips,
mean of pedestals looks generally uniform within the same chip.

## Slide 16: Pedestal Homogeneity: Width
Sigma of Gaussian
SCA = 0 (Memory-cell dependence is referred later.)
P3-chip0 looks strange.
Width of pedestal is almost uniform (3~4) throughout.

## Slide 17: Pedestal Stability
Pedestal stability is confirmed in this beam time.
Temperature: P1
Pedestal mean: P1 chip0

## Slide 18: MIP event
MIP program is performed for mainly energy calibration of all the pixels.
Hit map: Sum of the triggered events
Event display: ADC output of single event
Electron energy: 3 GeV
ADC
P3
(320 µm)

## Slide 19: TDC with MIP
Correlation of TDC between slab 1 and 2
Select 1 ch (at the center of the beam), 450 < ADC < 500 (to avoid timewalk)
~10 / 1 ns at the normal slope: timing resolution ~ a few ns?
TDC calibration in progress.
Preliminary

## Slide 20: Correction of Time Walk
Time Walk: TDC depends on ADC.
ADC
TDC interval
before
after
timing resolution: 3.2 ns
Preliminary

## Slide 21: MIP spectrum
slab | P1 | P2 | P3 | K1 | K2
thickness | 650µm | 650µm | 320µm | 650µm | 650µm
MPV | 146.5 | 144.9 | 71.3 | 141.4 | 146.1
Ped_width | 3.0 | 3.0 | 3.3 | 2.8 | 3.1
S/N | 49.0 | 48.9 | 21.7 | 50.2 | 47.5
P1
P2
P3
K1
K2
MIP spectrums of each slabs are shown.
Pedestal is subtracted.
Fitted by Lan-Gaus function.
  Convolution of Landau × Gaussian
MPV: Most Probable Value
Definition of S/NADC:

## Slide 22: MIP calibration: Summary
MPV maps:
Higher statistics
Problems
P1
P2
P3
K1
K2

## Slide 23: MIP calibration
slab: P1
Preliminary
Chip0 is broken?
MPV
Error of MPV
χ2/ndf
entries
MPV

## Slide 24: MIP calibration
slab: P2
Preliminary
Low-statistics because
DIF was broken in position scan
MPV
Error of MPV
χ2/ndf
entries
MPV

## Slide 25: MIP calibration
slab: P3
Preliminary
Chip0 is broken?
MPV
Error of MPV
χ2/ndf
entries
MPV

## Slide 26: MIP calibration
slab: K1
Preliminary
HV connection is bad?
MPV
Error of MPV
χ2/ndf
entries
MPV

## Slide 27: MIP calibration
slab: K2
Preliminary
Good!
MPV
Error of MPV
χ2/ndf
entries
MPV

## Slide 28: MIP calibration: Summary
MPV maps:
Higher statistics
Problems
P1
P2
P3
K1
K2

## Slide 29: Shower event
TBD
Event display of shower event
Preliminary

## Slide 30: Hit Energy
Hit energy after MIP calibration
Preliminary

## Slide 31: Simulation with DDSim in iLCSoft
  We performed detector simulation for this beam test.
  FEV13:
    CF: 0.6mm
    Electronics(Air)
    PCB: K1: 1.6mm, others: 1.8mm
    Glue(Air): 0.08mm
    Si: 0.32mm or 0.65mm
    Glue(Air): 0.08mm
    Kapton(Cu): 0.06mm
    CF: 0.6mm
    Plastic(polyethylene): 5mm
Preliminary

## Slide 32
Remaining Issues

## Slide 33: Double pedestal / Retrigger
run 32015, slab P1
Mean
Width
All
Good
Retrigger
Hip map
Mean difference
entry ratio
Double pedestal
by retrigger
Retrigger event:
BCID is consecutive.
(badbcid≠0)

## Slide 34: Pedestal difference between ADC/TDC mode
We found the difference of pedestals between ADC/TDC mode.
Memory-cell dependence is not same.
In the first memory cell, the difference of typical Ped_mean is ~15.
Mean of pedestal
with ADC mode
Mean of pedestal
with TDC mode

## Slide 35: Pedestal difference between ADC/TDC mode
We found the difference of pedestals between ADC/TDC mode.
Memory-cell dependence is not same.
In TDC mode, SCA~2 is worse.
Width of pedestal
with ADC mode
Width of pedestal
with TDC mode

## Slide 36: Pedestal difference between ADC/TDC mode
We found the difference of pedestals between ADC/TDC mode.
Memory-cell dependence is not same.
In TDC mode, SCA~2 is worse.
Width of pedestal
with TDC mode
There are double pedestal even after bcid selection in TDC mode.
Work in progress...

## Slide 37: Summary
FEV13: P1, P2, P3, K1, K2 in Kyushu
BT 2019 DESY: 5 slabs fully working finally
Pedestal study
  Homogeneity and Stability is verified.
  In TDC mode, pedestals become worse probably because of retriggers.
MIP study
  MIP calibration is almost completed.
  S/N is obtained for 5 slabs
TDC study
  Time walk is corrected.
  Timing resolution is obtained, however we need more detail study using injection.
Shower study
  Simulation setup is in progress.
  Event building is done and shower event is confirmed.
slab | P1 | P2 | P3 | K1 | K2
thickness | 650µm | 650µm | 320µm | 650µm | 650µm
S/N | 49.0 | 48.9 | 21.7 | 50.2 | 47.5
Preliminary

## Slide 38
backup

## Slide 39: Simulation outputs
Hit energy distribution in each slabs
w/ Tungsten
w/o Tungsten
(320µm)
Preliminary

## Slide 40: SKIROC2A
64ch / chip × 16 chip / slab
Self trigger: Individual threshold control is available.
3 types of output: ADC with high/low gain & TDC
3 types of DAQ mode:
  1. ADC high & ADC low 2. TDC & ADC [high or low] 3. TDC & Auto Gain
Preamp
TDC
Slow Shaper
Fast Shaper
Low Gain
High Gain

## Slide 41: Pedestal difference between ADC/TDC mode
Reme

## Slide 42: FEV13: SiW-ECAL technological prototype

## Slide 43: Hardware update
FEV13
SMB
DIF
Previous problems
  Carbon frame was not optimized for FEV13.
  HV connection between SMB and flex was fragile.
Update: New carbon frame
Cut
HV socket

## Slide 44: Hardware update
Previous problems
  Carbon frame was not optimized for FEV13.
  HV connection between SMB and flex was fragile.
Update: Conductive adhesion

## Slide 45: FEV13
We used 5 FEV13 slabs.
There are 2 types of Silicon thickness.
Slab position against e- beam: →
P1 | P2 | P3 | K1 | K2
dif_1_1_1 | dif_1_1_2 | dif_1_1_3 | dif_1_1_4 | dif_1_1_5
650 µm | 650 µm | 320 µm | 650 µm | 650 µm
↑ beam direction is
from front to back

## Slide 46: Data Summary
All the run information is summarized on
  https://drive.google.com/file/d/1uQojIu9KlS9badhVrBf1LRFNt-kz62vV/view?usp=sharing .
  not perfect, could be improved.
June 27: DIF on P2 had broken down → June 28: replaced DIF and recovered
June 28: made script to record temperature & start measurement
July 1: Data transition from P1 was sometimes lost because of bad connection of HDMI

## Slide 47: Pedestal Analysis
We generated pedestal maps for all runs.
run_30003-006
_dif_1_1_1 (P1)