## Slide 1: Pixel TPC testbeam results
Yevgen Bilevych, Klaus Desch, Harry van der Graaf, Fred Hartjes, Jochen Kaminski, Peter Kluit, Naomi van der Kolk,
Cornelis Ligtenberg,
Gerhard Raven, and
Jan Timmermans
LCTPC January 2022
8 Quad Module

## Slide 2: Pixel TPC
(TimePix1) TPX3 chip Quad Module TPC plane
(2007-14) 2017 2018 2019
(Octopuce)

## Slide 3: 8-QUAD module with field cage
in red guard wires

## Slide 4
DESY testbeam June 2021
Mounting the 8 quad module between the silicon planes
sliding it into the 1 T PCMAG solenoid

## Slide 5
DESY testbeam June 2021

## Slide 6
DESY testbeam June 2021

## Slide 7
DESY testbeam Analysis
High statistics data taken with B=0, 0.5 and 1 T fields
Electron Beam momenta of 6, 5 and smaller sets with 4,3,1 GeV
The stager allowed to move the 8 Quad Module
  3 positions x and 4 in z (drift)
Some data was taken with a rotated Module to allow studies of e.g. ExB deformations
The Mimosa Silicon telescope was described and aligned using the corryvreckan software with the General Broken Lines (GBL) track model. The ‘corryv’ software was updated to allow for a curved track fit

## Slide 8: Setup with Telescope planes
2 1 0
beam exit beam entry
 z axis z=0
5 4 3

## Slide 9
DESY Telescope Alignment
B=0 T run 6905
Preliminary

## Slide 10
DESY Telescope Alignment
B=0 T run 6905
Preliminary

## Slide 11
DESY testbeam Module Analysis
Preliminary results for Run 6916 B=0 T p =6 GeV will be presented
Tracks are preselected using the Mimosa Telescope
Matching between Telescope track and hits local x 2 mm and z 5 mm
Performed an alignment of the 8 quad module:
  local frame xy and z (drift) positions; angles dx/dy and dz/dy
  Adjusted t0 and x per chip
Time slewing applied using ToT:
  dt (ns)= 18.6/ (0.1577+ToT(ms))
Drift velocity (E=280 V/cm) 78 mm/ns
A local fit is performed to the track hits using errors:
  sxy =250 mm and sz =425 mm (slight dependence on ToT)

## Slide 12
DESY testbeam Module Analysis
z global
x global
z global
y global
y local
y local
x local
z local drift
Track 1050 hits
c2xy = 912/1048
c2z = 1740/1051
(no asymmetric tail or outlier removal applied)
Preliminary

## Slide 13
DESY testbeam Module Analysis
sxy =250 mm and sz =425 mm 1M hits
Run 6916 B=0 T p =6 GeV
<mean> 6.4 mm
Preliminary

## Slide 14
DESY testbeam Module Analysis
Track selection:
more than 100 hits in concentrator 0 and 1
more than 50% of the hits shouldbe on the tracks
1060 selected tracks
Impressive 900 hits / track
module or tracklength 157.96 mm
Run 6916 B=0 T p =6 GeV
Preliminary

## Slide 15
DESY testbeam Module Analysis
Tracking precision:
position 9 mm (xy) 13 mm (z)
angle 0.19 mrad (dx/dy) 0.25 mrad
module tracklength = 157.96 mm
Note that in a B field because of the reduced diffusion the tracking precision will improve substantially
Run 6916 B=0 T p =6 GeV
Preliminary

## Slide 16: Pixel TPC conclusions
First preliminary results of the 8 Quad Module in the DESY test beam in June 2021 have been presented
The run 6916 B=0 T with p=6 GeV has been analysed
The Mimosa telescope has been aligned using the corryvrecan software and tracks fitted with the GBL package
The 8 quad module data is decoded and matched to the tracks
The single electron resolution is:
  sxy =250 mm and sz =425 mm (mean drift distance of 6.4 mm)
In total 1060 tracks were selected with 900 hits on track
The tracking precision: position 9 mm (xy) 13 mm (z) in angle 0.19 mrad (dx/dy) 0.25 mrad
Note that the module or tracklength is 157.96 mm
This is very promising: more precise results can be extracted

## Slide 17: Backup General Pixel TPC conclusions
A single chip GridPix detector was reliably operated in a test beam in 2017
  Single electron detection => the resolution is primarily limited by diffusion
  Systematic uncertainties are low: < 10 µm in the pixel xy plane
  dE/dx resolution for a 1 m track is 4.1%
  A Quad detector was designed and the results from the 2018 test beam presented
  Small edge deformations at the boundary between two chips are observed
    added guard wires to the module to obtain a homogeneous field
    After correcting the edges, deformations in the transverse plane shown to be < 15 µm
An 8-Quad module has been designed with guard wires
  Deformations in the transverse plane for one quad are shown to be < 15 µm
Test beam data taken at DESY in 2021 are being analysed
A pixel TPC has become a realistic viable option for experiments
  High precision tracking in the transverse and longitudinal planes, dE/dx by electron and cluster counting, excellent two track resolution, digital readout that can deal with high rates
  A double grid will allow to reduce the Ion back flow distortions substantially