CALICE SiW ECAL – Beam test preparation
Roman Pöschl
On behalf of the SiW ECAL Groups in CALICE:
        SiW Ecal Meeting – 27/01/2022

---
Status of prototype
DESY beam test: 7/3/22 – 21/3/22
A priori prototype is operational but ...
Running in November with 9.6X0 => Not enough to absorb electron showers
Will exchange 7 2.1mm W plates by 2.8mm W plates and equip 8 layers with 4.2mm W plates
7x0.8+8x1.2X0 = 15.2X0
2.8mm W plates of 20x20cm2 from LLR (+ one 5.6mm W plate)
Two more COBs completed before and after Xmas
Gluing at LPNHE in december
Commissioning ongoing
Stack runs currently with two FEV13 included
Integration work by Jihane and Yuichi

---
Towards beam test – To do list 27/1/22
Mechanics
2.8mm plates need to be cut to right dimensions
Alexandre, Alice, Evelyne care about it ->
News: Seems to be possible at LLR with water jet cutting, delay?
Tests at LLR this week
Cost estimation from ATS company
Production of rails (4 for 4.2mm W plates and 7 for 2.8mm W plates)
Items above are urgent and need to be settled this week!!!
COBs
Commissioning -> Passports
Test with new termination cards
Slab 29 doesn't look bad, Slab 33 seems to be noisier
Both have problems in Row 1 and 3 and in particular with Chip 8 (need that on Slab 32 Chip 8 couldn't be started
(Mechanical) integration into stack
FEV13
Commissioning -> Passports
Operation voltage of stack?
Should be settled while FEV13 are in stack
System
Need to decide which ones to replace
Need to think about optimal arrangement of slabs

---
Towards beam test – To do list 27/1/22
Requires further understanding
These rows are connected to AVDD

---
Towards beam test – To do list 19/1/22 cont'd
Data acquisition
Integration with AHCAL
A priori concept exists
Meeting with AHCAL people at the beginning of February
Integration with beam telescope -> No priority for March
Data monitoring and DQ checks
We still have neither a high level data monitoring nor an event display
(As far as I understand) it would be helpful if data are accessible to EUDAQ2
Shared memory, nfs (or otherwise) mounted disk
Date will be sent via TCP socket (Test during Adrian's visit?)
Data management
Data can be buffered on 1 TB local disk and then shipped to cernbox (as last time)
Roman will try to mount disk array onto DAQ computer during October
Several TByte for buffer, maybe useful for local DQ checks
Still open from last TB (shame on me)
Data storage on grid?
What else?

---
DESY Test beam – On site
Unfortunately, the “rule of eight” will still apply at DESY
The teams
Week 7/3/22 - 13/3/22
Adrian, Yuichi, Roman, Jihane, Dominique + 1 persons from AHCAL + 2 persons from SiECAL
Jiri will be at DESY but will not be vaccinated -> Part of combination work outside of beam area
Week 14/3/22 - 20/3/22
Will be entire Ecal week
Adrian (limited availability?), Yuichi, engineers for W setup
Roman could stay but could also step back to free place
Survey on availability will be launched today (finally)
Conditions (same as last time)
Participants need have PCR test not older than 48h and be vaccinated (including booster)
Currently only two shifters are allowed in the beam test hall, more during setup
Propose to run two shifts per day 8h-16h, 16h-0h. night shifts should be taken into account to account for time loss
during setup of common running
Coordination (proposal)
Overall responsibility: Roman (1st Week), ? (2nd week)
Run coordination: Adrian (at least first week, depends also on requirements from common running)

---
DESY Test beam – Run plan
1st week
Setup including setup of/with AHCAL (count two days)
Quick commissioning w/o tungsten
DAQ Integration with AHCAL
Further commissioning w/o tungsten and check of common data taking
Hickups may lead to interruptions delays of program
2nd week tungsten program
FEV13 tests?
Several energy points with high statistics
Different positions, e.g. Scan over wafer boundaries

---
Backup